# SCOPE:  This procedure applies to UPMC-Hanover Laboratory

# PURPOSE:

The Aeroscout MobileView Temperature system is a wireless monitoring system that will alert the user when a device being monitored is out of specification, so that the product integrity inside the unit being monitored can be maintained.  In order to verify system functionality, a Daily report is sent from MVnotifications@upmc.edu @0830AM.  This report is reviewed daily by lab technical staff.

REAGENTS/SUPPLIES:

- MobileView IT HAN Lab- Daily Temps report sent via email to all users on the

HAN-LABMOBILEVIEW  email distribution list.

PROCEDURE:  Daily Temperature Report

| Step | Action |
| --- | --- |
| 1. | From OUTLOOK - UPMC email Open “HAN-Lab- Daily Temps” report- by preview |
| 2. | Report looks like this: |
| 3. | Scroll through the list- verifying all of the temperature are colored green. - If not- check to see the issue, document outliers in MobileView action alerts. - Will not return to green until temperature returns to normal range. - Continue to monitor unit’s temperature and process if critical range alert occurs. |
| 4. | At bottom of list- verify the number of units reporting are the current number of units (tags) in use in the laboratory. (48 is normal – currently 46 since 2 units out of service in microbiology) - If number does not correlate- then farther investigation is needed to see what unit is missing and why- document on Functionality Record Log. |
| 5. | Complete form MobileView Functionality Record- located on clipboard in Urinalysis Department. |

Document History

| Date of Origination and Document Control Number | January 1 2025 | MobileView replaces VersaTrak electronic temperature monitory system. 10/2023 Jan 1, 2015 |
| --- | --- | --- |
|  |  | Prepared by: Marilee Klunk, MT(ASCP). |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |

| Keywords: Mobileview Temperature |  |
| --- | --- |