# SCOPE: This procedure applies to UPMC-Hanover Laboratory

# PURPOSE:

The Aeroscout MobileView Temperature system is a wireless monitoring system that will alert the user when a device being monitored is out of specification, so that the product integrity inside the unit being monitored can be maintained. In order to verify system functionality, a Daily report is sent from MVnotifications@upmc.edu @0830AM. This report is reviewed daily by lab technical staff.

**REAGENTS/SUPPLIES:**

* MobileView IT HAN Lab- Daily Temps report sent via email to all users on the

HAN-LABMOBILEVIEW email distribution list.

**PROCEDURE: Daily Temperature Report**

|  |  |
| --- | --- |
| Step | Action |
| **1.** | **From OUTLOOK - UPMC email**  **Open “HAN-Lab- Daily Temps” report- by preview** |
| **2.** | **Report looks like this:** |
| **3.** | **Scroll through the list- verifying all of the temperature are colored green.**   * If not- check to see the issue, document outliers in MobileView action alerts. * Will not return to green until temperature returns to normal range. * Continue to monitor unit’s temperature and process if critical range alert occurs. |
| **4.** | At bottom of list- verify the number of units reporting are the current number of units (tags) in use in the laboratory. (48 is normal – currently 46 since 2 units out of service in microbiology)   * If number does not correlate- then farther investigation is needed to see what unit is missing and why- document on Functionality Record Log. |
| **5.** | Complete form MobileView Functionality Record- located on clipboard in Urinalysis Department. |

**Document History**

|  |  |  |
| --- | --- | --- |
| Date of Origination and Document Control Number | January 1 2025 | MobileView replaces VersaTrak electronic temperature monitory system. 10/2023  Jan 1, 2015 |
| Prepared by: Marilee Klunk, MT(ASCP). |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |
| Revision History/ Biennial Review: |  |  |

|  |  |
| --- | --- |
| **Keywords:**  *Mobileview Temperature* |  |
|  |