PRINCIPLE

The CoaguChek XS and the CoaguChek XS PLUS is intended for use by professional healthcare providers for quantitative Prothrombin Time (PT) testing for the monitoring of warfarin (Coumadin) therapy. The clotting factors II, VII, IX, and X in the coagulation cascade require activation via Vitamin K to form a blood clot and warfarin acts as an antagonist to Vitamin K, therefore preventing factor activation resulting in a reduced clotting effect. PT is measured and converted to units called International Normalized Ratio (INR). To minimize interlaboratory variability from formula variations in the thromboplastin, the following mathematical formula is used: (INR = PT patient/PT normal)ISI. A target therapeutic INR range is assigned to each patient based on their clinical diagnosis and risk factors. Patients’ INR are measured and their warfarin is managed through dosage adjustments to maintain their target INR to prevent unwanted blood clots (low INR), or bleeding (high INR). Coaguchek system calculates a patient’s INR by using fresh capillary or non-anticoagulated venous whole blood as opposed to plasma and is often used in an ambulatory setting for point-of-care (POC) results.

REAGENTS & EQUIPMENT

- CoaguChek XS Meter or CoaguChek XS PLUS Meter

- CoaguChek XS PT test strips

- Code chip

- Single use self-retracting Lancets

- Alcohol Pad (70% Isopropyl Alcohol)

- Gauze

- Band-aid

- Proper personal protection, including gloves

- Syringe, needle, tourniquet, capillary tubes, bulb if applicable

CoaguChek XS PLUS only

- CoaguChek XS Pro PT Controls: Level 1 and 2 with matching quality control chip, Diluent droppers, Scissors

STORAGE & HANDLING

Meter

  - Must be operated between 15o – 32oC (59°-89.6°F), and a relative humidity between 10% - 85%, without condensation.

  - Do not store meter in damp or humid conditions.

  - Do not touch any buttons while a test is in progress.

  - When testing, keep meter level.

  - Do not use the meter near a strong magnetic field such as a microwave oven, as this may interfere with testing.

  - The Coaguchek Meters automatically shut off if no buttons have been pushed. (Coaguchek XS-3 minutes, Coaguchek XS PLUS- 5 minutes

Test Strips and Code Chips

- Test strips must be store between 2o – 30oC (35.6°-86°F).

- If not using an electronic temperature monitoring system, storage temperature results must be documented on the temperature log once each day the clinic is operating. Monitoring of all non-working day(s) will take place on the first clinic day following the non-clinic day(s) by documenting the thermometer temperature.

- Test strips must be stored tightly capped in the original vial. Replace the vial cap immediately after removal of a test strip.

- Test strips must be used within 10 minutes of removal from the vial.

- Test strips are stable until expiration date on the vial. Do not use after expiration date, discard.

- Before using a test strip, ensure the code chip matches the code on the vial of test strips.

- Do not touch or remove a test strip while meter is performing a test.

- Do not remove or insert a code chip while meter is performing a test.

- Never open a vial or touch test strips with wet hands.

- Do not wait more than 15 seconds after sticking fingertip before application of blood.

- Do not add more blood after test has started.

Liquid Quality Controls (Only for the CoaguChek-XS PLUS Meter)

- Store controls in the refrigerator at 2o – 8oC (36o – 46oF).  DO NOT FREEZE.

- Unopened, lyophilized controls that are stored in the refrigerator are good until the expiration date.

- Discard any outdated controls.

- Controls are stable for 30 minutes after adding diluent.

METER SET-UP

Batteries

  - CoaguChekXS meter uses four AAA batteries. The Coaguchek XS PLUS meter uses four AA batteries.

  - When meter is turned on, display shows a battery symbol.

  - A fresh battery shows battery symbol with all four segments shaded.

  - Batteries must be replaced when only one segment is shaded in battery symbol.

  - If you store the meter for a period of time, it is recommended to remove the batteries.

Changing Batteries

  - Turn CoaguChek meter off by pressing the button with (I) on it.

  - Turn meter over and open battery cover by pressing latch gently inward while lifting.

  - Remove old batteries.

  - Place new batteries in battery compartment following diagram inside compartment.

  - Replace battery cover by snapping it into place.

Time and Date

  - Time and date must be set any time they are incorrect.

  - It is important that the time and date is correct as the meter compares current date to the test strip expiration date.

  - When replacing the batteries, to keep the date and time settings, the batteries must be inserted within 1 minutes for the Coaguchek XS, and 5 minutes for the Coaguchek XS PLUS.

Setting Time and Date for the CoaguChek XS

- Go to Setup Mode by pressing SET button located on left side of meter.

- The date format should be flashing in upper right-hand corner.

- Press SET button until year flashing; press M button to change or correct year; press SET button again to save.

- Press M button to change month; press SET button to save.

- Press M button to change day; press SET button to save.

- Time will flash in upper left-hand corner.

- Press SET button until hour flashes; press M button to change hour; press SET button to save.

- Press M button to change minutes; press SET button to save.

- Turn meter off by pressing (I) button.

Setting Time and Date for the CoaguChek XS PLUS

- From the Main Menu, touch Setup to open the meter settings.

- From the Setup Menu, touch Screen.

- From the Screen menu, touch Date/Time.

- From the Date/Time menu, touch Set Date.

- Use the ↑ or ↓ to select year, month, and day.

- Touch √ to save the setting or use X to delete.

- The display automatically returns to the previous screen.

- Touch Set Time.

- Use the ↑ or ↓ to select hour and minute.

- Touch √ to save the setting or use X to delete.

CODING (CALIBRATION)

- The CoaguChek XS and Coaguchek XS PLUS meter must be calibrated with a code chip:

  - Before meter is used for first time.

  - Whenever a new lot number of test strips are used.

  - Whenever a code chip error is displayed.

NOTE: It is IMPORTANT that the code chip matches the code on the vial of test strips. The code chip provides the meter with information such as lot number and expiration date of test strips.  A capital S in front of the number indicates that this code chip is for test strips. (A capital C in front of the number indicates that it is a control solution code chip included in each box of external control.)

Loading Code Chip

- Collect the meter and the test strips with code chip.

- Remove code chip from test strip box.

- Turn meter off by pressing power button (designated with (I) on it).

- Remove the old code chip by pulling straight out. The code chip is located on the left-hand side of the meter below the SET button.

- Insert new code chip with numbers facing upward and snap into place.

- Turn meter on to verify 3-digit code on screen matches 3-digit code on vial of test strips.

- Leave code chip in meter when not in use to maintain clean electrodes.

QUALITY CONTROL

Internal Controls

- Both meters automatically perform a two-level, on-board quality control test on the test strip before it displays each patient test result. “QC” appears in the display. A checkmark appears after the “QC√”, if the QC is successful.

- Internal controls check for accuracy and precision of analytical process.

- Internal controls check the following:

  - Electronic components and functions

  - Temperature of the test strips while test is in progress. (It is important to have reagents and test strips shipped and stored properly).

  - Expiration date and lot information on test strip.

  - On-board quality control tests in the same chamber that the patient is analyzed.

Documentation of Internal QC

  - Internal quality control results must be documented at least once daily. If acceptable, patient tests results can be reported.

  - All internal quality control results, in addition to any corrective actions taken to fix unacceptable control results, will be documented as “passed” (√) or “failed” on the patient log at the site or in procedural notes in the Standing Stone anticoagulation software program used in the Coumadin Clinics.

  - A “pass” result means that the quality control process met the criteria of acceptability, and that patient testing can continue. A “fail” result on the quality control would need remedial action taken prior to patient testing.

  - The lot number of test strips must be documented on a log each clinic day in the Coumadin clinics. The documentation must include the date, test strip lot number, expiration date, and the operator initials. A similar patient log to also include patient name, date of birth (DOB), patient ID (if required), and results of the internal quality control, will be used at other sites for each new lot number of test strips.

  - Internal quality controls will be documented in the HIS.

  - Internal quality control logs must be retained in accordance with the McLaren Healthcare Record Retention Schedule.

  - Quality control and patient logs will be reviewed for by the appropriate authority for potential problems and completeness of documentation.

External Liquid Controls (Optional for CoaguChek XS PLUS)

Roche Diagnostics does not provide liquid external quality control material for the CoaguChek XS with this test system. There are optional liquid external controls for the CoaguChek XS PLUS meter.

*Liquid Control Preparation

            - Insert quality code chip into the meter. This tells the meter the acceptable ranges for this box of controls.

            - Open lid of bottle and remove the rubber cap.

            - Hold the sealed dropper with the neck pointing upward, then cut off the end of the cap with scissors. Do not hold it near your face.

            - Apply gentle pressure to the reservoir and transfer the entire contents of the dropper to the bottle. Make sure the dropper does not come in contact with the dried control plasma.

            - Close the bottle.

            - Save the dropper for the next step.

            - Swirl the bottle using a circular motion to dissolve the control plasma. DO NOT SHAKE OR TURN THE BOTTLE UPSIDE DOWN.

            - Let the bottle sit undisturbed for 1 minute.

            - Control solution is now ready to apply to the test strip.

            - Use the reconstituted quality control within 30 minutes.

*Liquid Control Testing Procedure

- Place the meter on a flat surface, free of vibrations. Or hold it in your hands so it is horizontal.  Do not move the meter during testing.

- When you are ready to test, remove one test strip from the container and immediately close the container. Make sure it seals tightly.

IMPORTANT: DO not open a container of test strips or touch a test strip with wet hands or gloves. This may damage the test strip

- Use the test strip within 10 minutes of removing it from the container.  Otherwise, you may get an error message and you will have to repeat the test.

- Hold the test strip so the lettering is facing upward. Slide the test strip into the test strip guide in the direction indicated by the arrows.

- Slide the test strip as far as you can into the meter.  This turns the meter on. A beep tone indicated the meter has detected the test strip.

- Touch CONTROL TEST.

- The meter automatically checks to see if you have the correct test strip code chip. The three-digit code on the test strip container must match the number on the test strip code chip.

- Select the code already stored for your current control solution or touch NEW CODE to use a new control solution.

Note: When you first run your control, the QC TEST screen will not display. This screen will display the next time you use the control.

- If you are using a new control solution, remove the code chip from the meter and insert the code chip that came with the control solution.

- Select level (L1 or L2).

- The hourglass symbol shows that the test strip is warming.

- The dropper symbol flashes to indicate that the meter is ready and is waiting for the control solution to be applied. A 180 second count down begins. You must apply sample during this time.

- Swirl the control bottle to mix.  DO NOT mix the control with the dropper.

- Draw control solution into the dropper and put one drop on the top of the target area (clear, semicircular area of the test strip). DO NOT add more control or touch the strip while the test is in progress.

- The flashing dropper symbol changes to an hourglass symbol and a beep sounds when the meter detects sufficient sample. The test starts.

- Results will display in about one minute. Results are automatically saved to the memory.

- The acceptable range is displayed below the result. If a control test fails, an up arrow (too high) or down arrow (too low) flashes on the screen.

- Record the results on the quality control log. Report any corrective actions taken if control failed.

- Remove the quality control code chip for the meter and store it with the open box of controls. Re-insert the test strip code chip.

- Turn the meter OFF.

Documentation of External QC

    - All quality control results including any corrective action are recorded on the quality control log. This log includes the test strip lot number and expirations date, control lot number and expiration date, meter serial number, control ranges, operator identification, and date the test was performed.

    - The liquid quality control log is reviewed monthly by the POCC or another designee.

    - External quality control logs must be retained in accordance with the McLaren Healthcare Record Retention Schedule.

SPECIMEN COLLECTION & HANDLING

  - Venous blood - samples can be analyzed using CoaguChek meter. Collect sample in a plastic syringe free of anticoagulant. The syringe needle should be 23 gauge or larger. Capillary tubes cannot be glass or contain anticoagulants.

  - Finger stick - specimens may be applied directly from the patient’s finger onto the CoaguChek test strip. It is an option to have the sample collected in CoaguChek capillary tubes/bulbs and subsequently applied to test strip.

  - Blood must be analyzed within 15 seconds of finger stick.

  - Do not use the CoaguChek System for testing patients being treated with any direct thrombin inhibitors. These include Angiomax (bivalurudin), Argatroban, and Dabigatran etexilate (Pradaxa).

  - Sample size CANNOT be less than 8µL.

  - Additional blood CANNOT be added to the test strip one testing has begun.

**NOTE: If patient sample is drawn and the CoaguChek system indicates an error- draw a new specimen, using a different finger.

PATIENT PREPARATION

- Follow proper universal precaution guidelines when collecting and preparing patient specimens; see individual subsidiary infection control guidelines.

- Explain the purpose of test and steps of procedure before testing.

- If possible, have patient wash hands prior to testing.

- Check the date and time are correct for the meter.

- Check test strip expiration date.

TESTING PROCEDURE FOR COAGUCHEK XS METER

  - Place meter on solid horizontal surface or hold horizontally, do not move meter while test is in progress.

  - Remove test strip from container and securely replace lid. The test strip must be used within 10 minutes from removal from vial.

  - Slide test strip into guide in the direction the arrow is pointing until it stops. Meter will turn on.

  - Check time/date and battery level on meter. If time/date or battery level is incorrect or low, fix the problem. Refer to “Set-up Batteries, Date and Time”.

  - Confirm that the code number displayed matches the code number on the test strip container, then press “M” button.

  - An hourglass appears while meter warms up, this takes about 30 seconds. When meter is warmed up, a flashing test strip appears, and it will start counting down from 180 to 0 seconds. Sample must be placed on test strip within this time.

  - Find target area on strip.

  - While meter is counting down, collect and apply sample.

    - Finger stick method: DO NOT wipe away first drop of blood. Apply blood drop to the semicircular, transparent sample application area on top of the test strip or you can touch the blood drop against the side of the application area until the flashing drop of blood disappears. This must be done within 15 seconds of puncture. (It is OK to pick up meter to patient’s finger and touch edge of test strip but be sure to keep meter flat and set down after sample is collected).

    - Capillary tube method: Touch the CoaguChek capillary tube to the first blood drop after the finger stick and allow tube to fill up halfway by capillary action. Hold capillary tube over the semicircular, transparent sample application area on top of the test strip and expel specimen within 15 seconds of puncture. Avoid air bubbles.

    - Venous sample: EXPEL FIRST FOUR DROPS of blood from syringe or syringe needle. After drops are expelled, immediately place sample on target area of test strip.

  - While placing blood on target area, hold there until flashing drop of blood symbol disappears.

**NOTE: Do not add more sample while test is processing.

- Place used test strip, lancet, alcohol, gauze and other material in appropriate waste containers

11.	Shut meter off by pressing (I) button.

TESTING PROCEDURE FOR COAGUCHEK XS PLUS METER

- Place meter on solid horizontal surface or hold horizontally, do not move meter while test is in progress.

- Remove test strip from container and securely replace lid. The test strip must be used within 10 minutes from removal from vial.

- Slide test strip into guide in the direction the arrow is pointing until it stops. Meter will turn on.

- Enter your operator ID (See CoaguChek XS Plus System User Manual.)

- Touch√ [OK] to log on and move to the main menu.

- Touch PATIENT TEST.

- Select or enter Patient ID.

- Touch √ [OK]. An hourglass symbol indicates the test is warming up.

- Confirm that the code number displayed matches the code number on the test strip container.

- An hourglass appears while meter warms up, this takes about 30 seconds. When meter is warmed up, a flashing test strip appears, and it will start counting down from 180 to 0 seconds. Sample must be placed on test strip within this time.

- Find target area on strip.

- While meter is counting down, collect and apply sample.

    - Finger stick method: DO NOT wipe away first drop of blood. Apply blood drop to the semicircular, transparent sample application area on top of the test strip or you can touch the blood drop against the side of the application area until the flashing drop of blood disappears. This must be done within 15 seconds of puncture. (It is OK to pick up meter to patient’s finger and touch edge of test strip but be sure to keep meter flat and set down after sample is collected).

    - Capillary tube method: Touch the CoaguChek capillary tube to the first blood drop after the finger stick and allow tube to fill up halfway by capillary action. Hold capillary tube over the semicircular, transparent sample application area on top of the test strip and expel specimen within 15 seconds of puncture. Avoid air bubbles.

    - Venous sample: EXPEL FIRST FOUR DROPS of blood from syringe or syringe needle. After drops are expelled, immediately place sample on target area of test strip.

- While placing blood on target area, hold there until flashing drop of blood symbol disappears.

**NOTE: Do not add more sample while test is processing.

- The internal “QC” appears in the display. A successful outcome of the quality control test shows a checkmark that appears after the “QC√”.

- Place used test strip, lancet, alcohol, gauze and other material in appropriate waste containers

- Shut meter off by pressing (I) button.

DOCUMENTATION OF PATIENT RESULTS

  - Document the PT/INR result on a patient test log and in the Patient Electronic Medical Record (EMR) according to each subsidiary guideline. The information on the patient log will include the date, patient name, DOB or Patient ID, lot number of test strips and expiration date, internal QC result, and the signature of the operator (with credentials).

  - All Coumadin clinic patient visits will be documented in the Standing Stone anticoagulation software, Coagclinic. The information will include, but not limited to; date, time, initials of the operator, patient name, patient ID, INR results, internal quality control result and progress notes including resulting dosing and follow-up appointment date.

  - CareNow clinic patient results will be accepted/rejected by the operator on the meter. Acceptable results are transmitted from the Coaguchek XS PLUS meter to the patient EMR by Telcor QML middleware.

  - The CoaguChek-XS meter automatically stores up to 300 test results and their dates and time in its memory. When viewing results in memory, results can be changed between INR and Protime seconds by pressing the Set button on the meter side. This allows the user to toggle between Protime seconds and INR.

  - The CoaguChek-XS PLUS meter stores 2000 patient results and 500 quality control tests to memory.

**NOTE: All test results stored in memory are not accompanied by the required two patient identifiers and should not be used for reporting previous performed patient INR results.

INTERPRETATION:

Reportable Range = INR 0.8 – 8.0

Normal Range INR= 0.9-1.1

Critical Call result: = INR’s  5.0

Therapeutic Range= 2.0-3.0

Mechanical Heart Valve Range = 2.5-3.5

Recurrent Systemic Embolism Range = 2.5-3.5

EXPECTED VALUES

  - Any patient whose INR is  5.0, will be confirmed from a repeat INR using a different finger.

  - Any patient whose INR results are “out of the reportable range” (INR 0.8-8.0), will display an error message. Refer to the error codes for reference in troubleshooting or contact Roche Diagnostics at 1-800-428-4674.

  - If a patient’s INR results using the CoaguChek system are 0.8 or 8.0, do not report the results. Send the patient to a laboratory for a STAT confirmatory venous draw.

  - Physicians will be notified of critical values on confirmatory venous draws within 4 hours by the performing laboratory.

  - Confirmatory testing with a venous draw should be performed anytime the pharmacist or provider has concerns regarding the patient’s INR trends or if clinically warranted.

  - Physician practices and clinics will follow the subsidiary Hospitals Critical Values Requiring Notification Policy for guidelines in the absence of specific medical-staff-approved policies.

  - All physician and providers recognize the INR differences between POC devices and laboratory instrumentation with INR’s 4.0. All clinical decisions will fall within the scope of the out-patient clinic in collaboration with the patient’s physician and/or Director of the clinic.

COAGUCHEK-XS ERROR MESSAGES

- Error: Test Strip – May occur when test strip inserted while meter turned on. May also occur if test strip not removed after performing test, or if meter timed out after strip inserted. May also mean test strip unusable or not a CoaguChek-XS PT test strip. Solution: Remove test strip and replace with new CoaguChek-XS PT test strip.

- Error: Meter Temperature – Occurs when meter is too warm or cold to give correct measurements. Solution: Turn meter off and store at temperatures between 15o – 32oC for about 30 minutes.

- Error: Battery – Battery level low. Solution: Replace batteries.

- Error: Test Strip Guide Cover – Test strip guide cover is not closed. Solution: Close test strip guide cover by snapping into place.

- Error: Code Chip – Code chip is missing, damaged, or improperly inserted. Solution: Check proper insertion of code chip. If code chip is damaged, call the Roche Diagnostics Technical Service Center at 1-800-428-4674.

- Error: Test Strip Expired (Error 3) – Test strips are beyond expiration date. Solution: Ensure meter’s date and time is correct. If correct, turn meter off and replace new code chip from new in-date test strips.

- Error: Test Strip Unusable (Error 4) – Test strip is unusable. Solution: Turn meter off, remove test strip, reinsert. If message still reads unusable, replace with new test strip.

- Error: Time Exceeded (Error 000) – Blood not applied to test strip within 3 minutes. Solution: Turn meter off, remove test strip. Repeat test using same test strip, but new drop of blood from different finger. If blood was added to reagent test strip, discard test strip and repeat tes

- Error: Blood Application (Error 5) – Error occurred while applying blood to test strip. Solution: Turn meter off and remove test strip, repeat test using new test strip and new sample from different finger.

- Error: Test Strip Interference (Error 6) – Test strip removed or touched during test. Solution: do not touch or remove test strip during tests, repeat test using new strip and a new sample from a different finger. May see “ERROR 6” with patients under treatment on warfarin in combination with antibiotics and/or chemotherapeutics (i.e. INR8.0).

- Error: Quality Control Failure (QC error) – Test strip quality control failed, test strip unusable. Solution: Turn meter off, remove test strip, repeat test using new strip and new sample from a different finger.

- Error: Measurement Error (Error 7) – Measurement error due to blood sample. Meter was      unable to detect a clot. Solution: Do not remove or touch test strip during testing. Apply blood within 15 seconds of finger stick. Sometimes error message occurs in patients with long coagulation times (INR 8.0). Repeat the test. If message appears again, result must be checked using alternative method. If using capillary tubes, make sure they are CoaguChek-XS capillary tubes.

- Error: Internal Error (Error 8 or 9) – Occurred during internal diagnostic tests. Solution: Turn meter off, remove batteries, and wait at least 1 minute. Reinsert batteries, reset the date and time. Repeat test, if same error message occurs, meter has a defect and may need to be replaced. Call the Roche Diagnostics Technical Service Center at 1-800-428-4674.

COAGUCHEK-XS PLUS ERROR MESSAGES

Status Messages

Code Text

| I-001 | The battery is getting low. Replace or recharge battery pack as soon as possible. |
| --- | --- |
| I-002 | The patient ID you entered is not on the list. |
| I-004 | Only 100 tests can be performed before the meter memory is full. Dock meter or disable this setting. |
| I-005 | Communicating with DMS. Please wait. |
| I-006 | Only 500 tests left before the maximum number of tests is reached. Contact your Roche representative. Only 100 tests left before the maximum number of tests is reached. Contact your Roche representative. |
| I-009 | A comment is required if you reject a result. Please enter a comment. |
| I-010 | Observer (USERNAME), please log in to proceed. Otherwise the OTS test will not be complete. |
| Error Messages |  |
| Code Text |  |
| E-101 | A patient ID is required. Enter a valid patient ID. |
| E-102 | The ID was not found on the list of valid patient IDs. Enter a valid patient ID. |
| E-103 | An operator ID is required. Enter a valid operator ID. |
| E-104 | The ID was not found on the list of valid operator IDs. Enter a valid operator ID or see your System Administrator. |
| E-105 | Invalid password. Enter valid password or see your System Administrator. |
| E-107 | There are no operators loaded in this meter. In order to continue, operators must be downloaded. Please see your System Administrator. |
| E-108 | The passwords you entered do not match. Please try again. |
| E-109 | The scanned barcode is invalid. Scan a valid barcode or see your System Administrator. |
| E-111 | The observer ID is invalid. The observer supervising the current OTS test must log in again. |
| E-201 | The strip lot with code <Number> has expired. Check the meter date/time setting and use a non-expired strip lot. |
| E-202 | The control lot has expired. Check the meter date/time setting and use a nonexpired control lot. |

| Error Messages |  |
| --- | --- |
| Code Text |  |
| E-203 | The lot with the code <Number> is not allowed on this meter. See your System Administrator. |
| E-204 | The barcode on the test strip cannot be read. Please check the barcode on the test strip and repeat the test. |
| E-205 | Strip removed during test. Repeat test, using a new test strip if necessary. |
| E-207 | Invalid code chip. Remove the code chip and repeat the test. |
| E-210 | A quality control test must be performed. To continue, perform a quality control test. |
| E-211 | The test strip guide cover was removed during a test. Repeat the test. |
| E-301 | The date is invalid. Enter the correct date and time. |
| E-302 | Battery low. No more tests can be performed. Replace or recharge battery pack. |
| E-303 | The maximum number of tests has been reached. Contact your Roche representative. |
| E-304 | Meter is too cold or too warm to test correctly. Power meter off and repeat test after 30 minutes at room temperature. |
| E-305 | The meter memory is full. Dock the meter or disable this setting. |
| E-401 | Sample detection has failed. Remove the test strip and repeat test, using a new test strip if necessary. |
| E-402 | A test strip error was detected. Check your procedure and repeat test, using a new test strip if necessary. |
| E-403 | A dosing error was detected. Check your procedure and repeat test, using a new test strip. |
| E-405 | A measurement error occurred. Check your procedure and repeat test, using a new test strip. |
| E-406 | Unsuitable sample composition. Check your procedure and repeat test with new strip and fresh sample. Check medication regimen of patient. If problem persists, check result with another method and refer to manual for more detail. |

| E-407 | Strip failed onboard control. Remove the strip and repeat the test, using a new strip from another test strip container if necessary. |
| --- | --- |
| E-501 | The self-check detected a general error. Repeat the self-check. If problem persists, power meter off and on again. |
| E-601 | Code chip cannot be read. Remove the code chip and repeat the test. |
| E-602 | General System Error. The meter is defective. No test is possible. Contact your Roche representative. |
| E-603 | Temperature or heater error. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-605 | Hardware or electronic error. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |

| Error Messages |  |
| --- | --- |
| Code Text |  |
| E-606 | Measurement sequence error. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-607 | Error with measurement module. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-608 | A software error was detected. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-701 | Failed to connect to DMS. Unlock meter from Base Unit and then redock it. If problem persists, power meter off and on. |
| E-702 | Communication error with DMS. Unlock meter from Base Unit and then redock it. If problem persists, power meter off and on |
| E-704 | ICI communication error with DMS. Power meter off and on, and repeat the last procedure. |
| E-705 | Communication error with measurement module. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-706 | Communication error with test strip barcode scanner. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-708 | General communication error. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-801 | Patient database cannot be read and will be deleted. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-802 | QC database cannot be read and will be deleted. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-803 | Comments database cannot be read and will be deleted. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-804 | Setup database cannot be read and will be deleted. Power meter off and on, and repeat the last procedure. If problem persists, contact your Roche representative. |
| E-805 | The meter detected a fault with the EEPROM memory. Power meter off and on. If problem persists, contact your Roche representative. |

| E-806 | The meter detected a fault with the Flash memory. Power meter off and on. If problem persists, contact your Roche representative. |
| --- | --- |
| E-807 | The meter detected a general fault with the memory. Power meter off and on. If problem persists, contact your Roche representative. |

MAINTENANCE

- Clean the meter with each patient testing. Clean/disinfect the CoaguChek meter housing for a contact time of  1 minute. Let dry for 10 minutes before performing testing.

- 70% isopropyl alcohol (70% isopropyl alcohol pad) is used to clean the meter between patient testings. Roche does not recommend using 10% bleach for cleaning or disinfecting the CoaguChek meter. Frequent use with 10% bleach may leave a residue within the system due to evaporation. This will result in conductive connections of the electrical contacts.

- Turn meter off during cleaning procedures.

- Always wipe meter with lint-free tissue.

Cleaning the test strip guide

- Clean test strip guide by opening cover. Use thumbnail to open test strip guide cover by pressing front edge upward. Clean cover.

- Holding meter upright with the test strip guide facing down. Clean test strip guide with cotton swab or pad using 70% isopropyl alcohol. Be sure the swab is only damp, not soaking wet, to ensure excess fluid does not enter the meter.

- Leave test strip guide cover off for 10 minutes to dry.

- Close cover by snapping it into place.

LIMITATIONS OF THE PROCEDURE

- The blood sample must be at least 8 µL.  Low sample volume will cause an error message to appear on the meter.

- The CoaguChek System uses only fresh capillary or non-anticoagulated venous whole blood. Plasma or serum cannot be used.

- Use only plastic syringes without anticoagulants or additives.  Do not use glass tubes or syringes.

- Never add more blood to test strip after test has begun or perform another test using the same finger stick.

- When a patient is on intravenous infusion therapy, do not collect sample from arm receiving the infusion line.

- No significant effect on test results was seen with:

- Bilirubin up to 30 mg/dL

- Lipemic samples containing up to 500 mg/dL of triglycerides

- Hemolysis up to 1000 mg/dL

- Hematocrit ranges between 25-55%

- Heparin concentrations up to 0.8 U/mL

- Clopidogrel up to 20 mg/dL

- Fondaparinux up to 5 mg/L

- Low molecular weight heparins (LMWH) up to 2 IU anti-factor Xa activity/Ml

NOTE:  The presence of anti-phospholipid antibodies (APAs) such as Lupus antibodies (LA) can potentially lead to prolonged clotting times, i.e., elevated INR values.  A comparison to an APA-insensitive laboratory method is recommended if the presence of APAs is known or suspected.

ELECTIVE GUIDELINES FOR MANAGING THE COAGUCHEK XS and COAGUCHEK XS PLUS METER

Anticoagulation clinics using CoaguChek-XS or CoaguChek-XS PLUS devices should be validated against their subsidiary laboratory mainframe instrument before being used for patient testing.

- Validation study- A validation study should be performed by the Anticoagulation clinics to demonstrate correlation between the CoaguChek meter and the Laboratory instrumentation. The validation should cover low, mid-range, and high INR patient values by using data collected on 20 paired INR samples (CoaguChek and Laboratory mainframe instrument).

- Collect one sample in a blue 3.2% citrate blood collection tube and send to the laboratory for the INR.

- Collect a second sample by finger-stick and perform an INR on the CoaguChek meter.

- Enter the data from both methods using a software program for method comparisons such as the EP Evaluator. The comparative method is the current method or method being compared to and is entered as the X value (lab instrument). The new method is entered as the Y value (CoaguChek).

- Lot number change/new shipment of strips - A Split Sample study should be performed by the anticoagulation clinics using 3 to 5 samples for testing the INR using the laboratory mainframe instrument and the CoaguChek meter for lot number change/new shipments of Coaguchek strips.

- Data from both the validation and lot number change/new shipments are reviewed and signed by the Laboratory Director or designee.

- Acceptable results are the recommended standard deviation range per manufacturer as defined by the Clinical laboratory. Results of the compared methods in the validation study and the lot number change/new shipment of strips should correlate within +/- 30%.

- Communicate with clinicians/staff on reasons for discrepant INR results (CoaguChek vs. Lab)

- Different lots of reagent, with different ISI and calibrations

- Different reagents and device/reagent combinations

- Different brands, and even lots, of blood collection tubes

- Alternating measurements by different methods (CoaguChek vs. Lab)

- Incorrect calculation of the INR and/or Mean Normal Protime range calculated with reagent lot number changes.

- Specimen type (whole blood vs. platelet poor plasma)

- Investigate discrepancies

- Patient specific findings - patient’s current clinical presentation

- Pre/Post analytic errors

- Method/reagent differences

- Lot shifts, Proficiency/ Quality Control performance history

- Proficiency Samples are offered through College of American Pathologists (CAP) and the American Proficiency Institute (API) quarterly and are recommended to be performed to ensure quality performance as indicated by a survey report.

TRAINING AND COMPETENCIES

Only personnel who have successfully completed the Competency and Training Program provided by a Laboratory Designee or approved trainers will be authorized to perform POC testing. Competencies will be retained at each site.

Initial training includes:

- Use appropriate training check list.

- Observe trainee performance.

- Complete assigned McLaren HealthStream module or Med-Training Module.

Annual Competencies includes:

- On-line modules and direct observation are repeated annually for competency assessment.

- When problems are identified with an individual's performance.

- The annual assessment of competency can be performed throughout the entire year to minimize impact on workload

- Interfaced Coaguchek XS meters use Auto-certification to recertify operator competency annually.

REFERENCES

CoaguChek- XS System User Manual, Roche Diagnostics, 2011

CoaguChek XS PLUS System Operators Manual, Roche Diagnostics, 2018