# VENTURA COUNTY MEDICAL CENTER

# SANTA PAULA HOSPITAL

#### Associated with the UCLA School of Medicine

| PROCEDURE Directigen Meningitis Combo Test |  | POLICY NO: Serology 3.4 | PAGE 1 OF 8 |
| --- | --- | --- | --- |
| DEPARTMENT: Laboratory AFFECTS: Microbiology | EFFECTIVE DATE: 4/5/1993 | REVIEW DATES: 5/1994,9/1995,10/1997,2/2000, 12/2001,12/2005,4/2006,3/2007, 1/2008,3/2009,3/2010,3/2011, 12/2015 |  |
|  |  | REVISION DATES: 12/2013 |  |
| INTENDED USE The Directigen Meningitis Combo Test is a presumptive latex agglutination test for the direct qualitative detection of antigens to Haemophilus influenzae type b, Streptococcus pneumoniae, Neisseria meningitidis groups A, B, C, Y or W135 and Escherichia coli K1 in cerebrospinal fluid (CSF), serum or urine. The test can also be used for the direct qualitative detection of antigens to Streptococcus agalactiae (group B) in CSF and serum. In addition, the test kit provides confirmation and serogrouping capabilities from suspected colonies of H. influenzae type b, S. pneumoniae, group B Streptococcus, and N. meningitidis groups A, B, C, Y or W135. Visible agglutination occurs when a sample containing any of these bacterial antigens is reacted with its respective antibody-coated latex beads. SUMMARY AND EXPLANATION The diagnosis of bacteremia and meningitis, especially in young children, can be difficult. As many as 55% of children are seen by a physician and started on antibiotics before meningitis is detected. Detection of microbial antigens in the CSF is a rapid and helpful method for diagnostic microbiology. It may be the single most important test in cases of partially treated meningitis since Gram stain and culture may be negative. Detection of a specific antigen is a clinically significant finding and a valuable aid when choosing antimicrobial therapy. Haemophilus influenzae type b, Neisseria meningitidis, and Streptococcus pneumoniae have been reported to be the three causative agents responsible for approximately 84% of cases of bacterial meningitis. Streptococcus agalactiae (Group B) and Escherichia coli KI are major bacterial pathogens in the newborn. Strains of group B streptococci and E. coil KI frequently colonize in the vagina and/or rectum and may be associated with maternal septicemia and neonatal septicemia, pneumonia and meningitis. The E. coli KI polysaccharide antigen has been shown to be structurally and immunologically similar to Neisseria meningitidis group B antigen. The Directigen Neisseria meningitidis group B Latex Reagent does not differentiate the two antigens, but can be useful in the diagnosis of neonatal E. coil KI meningitis. The high morbidity and mortality associated with Group B streptococci and E. coli KI in newborns make rapid, accurate identification of these organisms extremely important. Immunological methods for detecting characteristic exoantigens of pathogenic microorganisms in patient fluids (CSF, serum, and urine) are typically faster than traditional methods such as culture. These techniques include counterimmunoelectrophoresis (CIE) and latex agglutination. The latex agglutination procedure has been found to be more rapid and sensitive than CIE in the detection of purified antigen. PRINCIPLES OF THE PROCEDURE Specific antibodies are bound to the surface of latex beads. Latex particle aggregation becomes large enough to allow rapid visualization of positive agglutination in the presence of specific antigens. These specific soluble polysaccharide antigens accumulate in CSF, serum or urine as a result of infection by H. influenzae type b, S. pneumoniae, and N. meningitidis groups A, B, C, Y or W135 and Escherichia coli Kl; and in CSF and/or serum, as a result of infection by Streptococcus agalactiae (group B). All of these antigens can be detected with the Directigen Meningitis Combo Test Kit. REAGENTS Directigen Meningitis Combo Kit: Rabbit Polyclonal antibody-Coated latex Suspensions of the following: Reagent 1- Anti-H. influenzae type b Reagent 2 - Anti-S. pneumoniae Reagent 3 - Anti-Group B Streptococcus Reagent 4 – Anti-N. meningitidis, Groups C and W135 Reagent 5 – Anti-N. meningitidis, Groups A and Y Mouse Monoclonal Antibody-Coated Latex suspension of the following: Reagent 6 – Anti-N. meningitidis, Group B/ E. coli K1 Rabbit Immunoglobulin-Coated Latex Suspension of the following: Reagent A and Reagent B Positive Control – Polyvalent Positive Antigen Control, H. influenzae type b, S. pneumoniae, group B Streptococcus and N, meningitidis groups A, B, C, Y and W135 antigens. Negative Control - glycine buffered saline Specimen Buffer Solution Warning: Reagents contain sodium azide, which is very toxic by inhalation, in contact with skin, and if swallowed. Contact with acids liberates very toxic gas. After contact with skin, wash immediately with plenty of water. Sodium azide may react with lead and copper plumbing to form highly explosive metal azides. On disposal, flush with a large volume of water to prevent azide build-up. . Test Cards: Cards must be flat for proper reactions. If necessary, flatten cards by bowing back in a direction opposite to that of the curl. Care should be taken not to finger-mark the test areas, since this may result in an oily deposit and improper test results. Use each card once and discard. Store cards in original package in a dry area at room temperature. When spreading within confines of test areas, avoid scratching card surface with the plastic stirrer. If the specimen does not spread to outer perimeter of test area, use another test area of the card. Rotation: The recommended speed for mechanical rotation is 100 ± 2 rpm, but rotation between 95 and 110 rpm does not significantly affect the results obtained. The rotator should circumscribe a circle approximately 2 cm in diameter in the horizontal plane. A moistened humidifying cover should be used to prevent drying of test specimens during rotation. Storage of Reagents: Upon receipt, remove carton containing reagents and refrigerate at 2 - 8°C. DO NOT FREEZE. Reagents should be recapped and returned to refrigeration when not in use, taking care not to mix color-coded caps. SPECIMEN COLLECTION AND HANDLING Refer to appropriate texts for details of specimen collection and handling procedures. Specimens should be tested as soon as possible; however, if the sample cannot be tested immediately, it should be stored at 2 - 8°C (for up to 48 h), or at -20°C. Serum must be separated from whole blood prior to testing or storage. Specimen Pretreatment The use of covered glass test tubes (borosilicate) in a 100°C (boiling) water bath is the most effective means of specimen treatment. Although glass test tubes in a heat block may be used, greater variation in the thorough heating of a sample may be noted as a result of variations in tube size and sample volume. - Specimen Preparation (CSF): - Heat specimens for 3 min at 100°C (e.g., water bath or heat block) and allow to cool to room temperature before use. For optimal sensitivity, N. meningitidis group B and E. coli K1 (Reagent 6) testing should be performed on unheated specimens (see "Limitations of the Procedure"). - For CSF specimens showing turbidity, centrifuge after heating for 10 min at 1400 x g prior to testing. The supernatant fluid is to be used as the test specimen. - Test specimens as described in "Procedures." - Specimen Preparation (Serum) 1. Dilute serum specimens of at least 0.6 mL 1:1 with Directigen Specimen Buffer and mix. - Heat specimens for 5 min at 100°C (e.g., water bath or heat block) and allow to cool to room temperature before use. - Using a wooden applicator stick, break up the protein "clot" formed, and vortex vigorously (approximately 5 sec). - Centrifuge at a minimum of 1400 x g for 15 min. - Test supernatant fluid as described in "Procedures." - Specimen Preparation (Unconcentrated Urine): 1. Dilute urine specimens of at least 0.4 mL 1:1 with Directigen Specimen Buffer and mix. - Heat specimens for 5 min at 100°C (e.g., water bath or heat block) and allow to cool to room temperature before use. - Centrifuge at a minimum of 1400 x g for 10 min. - Test supernatant as described in "Procedures." - Specimen Preparation (Concentrated Urine): - Urine samples that are turbid or have particulate material should be centrifuged at 1400 x g for 10 min before concentrating. - Urine samples may be concentrated 25-fold with a Minicon B-15 concentrator (Amicon Corporation, Danvers, Massachusetts). - Dilute at least 200 mL of urine concentrate 1:1 with Directigen Specimen Buffer and mix. - Heat specimens for 5 min at 100°C (e.g., water bath or heat block) and allow to cool at room temperature before use. - Centrifuge at a minimum of 1400 x g for 10 min. - Test supernatant fluid as described in "Procedures." NOTE: Urine cannot be used to test against group B Streptococcus (Reagent 3) - Preparation for Confirmation of Colonies from Culture: - Locate suspected colonies on the agar surface from 18 - 24 h cultures that meet morphological and Gram stain characteristics of organisms that are appropriate for testing with Directigen Meningitis latex reagents. A Gram stain should be performed prior to testing to ensure that organisms are appropriate for testing with the Directigen Meningitis latex reagents. - Pipette 0.5 mL, (approximately 10 drops) of Control - reagent into a small glass test tube (10 x 75 mm or equivalent). - Select several (2 - 3) isolated colonies of similar morphology from the original or subculture plate using a sterile loop and suspend into the above tube to achieve a suspension equal to a McFarland #1 turbidity standard. Note: Over inoculating will yield an excessively heavy suspension, which may cause erroneous results. - Heat the suspension for 3 min. at 100°C (e.g., boiling water bath or heat block) and allow to cool to room temperature before testing. - Centrifuge at a minimum of 1400 x g for 10 min. - Test supernatant as described under "Procedures." Note: If atypical agglutination patterns are observed, refer to "Limitations of the Procedure." PROCEDURES The testing area, reagents, test specimens and test components should be at room temperature (15-30°C) when used. Materials Provided: All materials as listed under "Reagents," work station, disposable test cards and accessories. Test Procedure (CSF, serum, urine, concentrated urine and colony confirmation): Remove reagents from refrigerated storage, and place reagents and card (or glass test slide) in designated wells of Work Station Tray. The testing area, reagents, test components and test specimens should be at room temperature (15-30 degrees C) before being used. Use only the test card or glass slide recommended for this kit. Before using the glass slide, thoroughly clean with a lint-free tissue. - .Dispense one drop of Control + onto circles 1 through 6 of row "+". Place one drop of Control -onto circles I through 6 of row "-." - Micropipette 50 µ1 of test sample onto circles 1 through 6, row S (Sample) and in circles labeled "A" and "B." Rows "+" (Positive) and "-" (Negative) are used for controls. Subsequent sample testing can be done using rows -I-, - and S of new (or cleaned) test cards (slides). - Holding the dispensing bottle by the cap, vigorously swing (without inverting) to thoroughly mix Reagents 1 - 6 and Reagents A and B. Before uncapping each bottle, gently tap base on counter top to assure no latex reagent remains in the tip. - Dispense one drop of Reagent A (Control Latex Suspension) onto the "A" circle. Repeat the procedure dispensing one drop of Reagent B (Control Latex Suspension) onto the "B" circle. - Dispense one drop of Latex Suspension Reagent 1 onto the circles in column 1, rows "+" (Positive), "-" (Negative) and "S" (Sample). Repeat the procedure for the remaining Latex Suspensions (Reagents 2 - 6) in rows "+", "-" and "S", columns 2 through 6. - Mix the samples and Latex reagents in each circle with a plastic stirrer, alternately using first one end, and then the opposite end for the next circle. Discard the stirrer. - Place the card or glass slide on a mechanical rotator and rotate at a speed of 100 ± 2 rpm for 10 min. Use a moistened humidifying cover to prevent evaporation. Immediately at the end of 10 min, read the test results macroscopically under a high intensity incandescent light. Note: For testing a specimen (CSF, serum, urine, concentrated urine or colony confirmation) with one individual Reagent 1 - 5, Reagent A must be used. For testing with one individual Reagent 6, Reagent B must he used. Use the above procedure. INTERPRETATION OF TEST RESULTS Record Control + and Control - test results first: The Control + should yield strong agglutination in circles of row "+" within 10 min. The Control - should show no agglutination in circles of row "-". Agglutination in any of the circles containing Control - renders the reaction uninterpretable. Do not use the kit if Control + and Control - do not yield appropriate results Record sample test results: Positive Test - Should show agglutination. Any degree of agglutination present in one of the latex reagents indicates the presence of the corresponding antigen. Agglutination in two or more latex reagents or the corresponding Reagent A (Reagents 1-5) or Reagent B (Reagent 6) renders the reaction uninterpretable. Negative Test - Should show no agglutination. Before reporting results of MAP, check to see if a bacterial culture is in progress. If no culture has been ordered, ask the physician to order a culture. REPORTING OF RESULTS Use the MAP worksheet to record the results. Record results of the kit lot number, expiration date, rotator speed and results of the quality control tests. Report results in LIS as positive or negative for each of the antigens tested. Place worksheet in the MAP binder. USER QUALITY CONTROL Include Control + and Control - testing with each batch of specimens tested as described in step 1, "Test Procedure." Quality control requirements must be performed in accordance with applicable local, state and/or federal regulations or accreditation requirements and your laboratory's standard Quality Control procedures. It is recommended that the user refer to pertinent NCCLS guidance and CLIA regulations for appropriate Quality control practices. Parallel Testing If a positive patient sample and/or a positive CAP specimen is not available for parallel testing, testing can be accomplished by using Streptococcus pneumonia ATCC 49619 and following the procedure for Confirmation of colonies from culture LIMITATIONS OF THE PROCEDURE These latex agglutination tests are not intended as a substitute for bacterial culture. CSF culture must be performed on both positive and negative CSF specimens. Diagnosis of bacterial meningitis infection is only possible with appropriate culture procedures. Samples with extremely low levels of antigen, for instance early in the course of infection, may yield false negative results. Furthermore, samples with exceedingly high antigen concentrations may exhibit prozone effects producing inappropriately negative results. Although not extensively studied, prozone phenomena have only been observed in specimens seeded with extremely high antigen levels, and not in clinical specimens. This assay is for the qualitative detection of antigens to H. influenza type b, S. pneumonia, N. meningitidis groups A, B, C, Y or W135 and E. coli K1 in CSF, serum or urine; and to group B Streptococcus in CSF and serum. In addition, the test kit can also be used with suspected colonies of H. influenzae type b, S. pneumoniae, group B Streptococcus and N. meningitidis groups A, B, C, Y or W135. Performance with other specimen types has not been assessed. The Directigen Meningitis Combo Test cannot be used with blood culture media. A positive or negative Group B streptococcal result is only indicative of the presence or absence of Group B streptococcal antigen and is not diagnostic for the presence or absence of group B streptococcal disease. This device should not be used as a substitute for bacterial culture in the diagnosis of group B streptococcal septicemia and/or meningitis. A positive or negative result is a presumptive result for group B streptococcal antigen. The result must be confirmed by culture. The only infant specimens recommended for the direct qualitative detection of group B streptococcal antigen are serum and cerebrospinal fluid. Testing infant urine with devices for direct qualitative detection of group B streptococcal antigen cannot be recommended at this time; currently there is insufficient performance data supporting the use of this test on urine as a reliable predictor of group B streptococcal disease. Pneumococcal and Haemophilus influenzae type b strains not possessing a capsular antigen may not be detected by immunological techniques. Cross-reactions of H. influenzae reagent are known to occur in the presence of Escherichia colt K100. N. meningitidis groups C and W135 reagent is known to cross react in the presence of Escherichia coli K92. Other cross-reactions, uninterpretables and false positives may occur which are reduced or eliminated by proper specimen preparation (see "Specimen Pretreatment" and "Specimen Preparation"). Studies with CSF specimens seeded with N. meningitidis group B and E. coil K1 antigen have suggested that some loss in sensitivity may result after heating as compared to unheated samples. Untreated serum and urines may yield uninterpretable results with the latex reagents and should be treated as in "Specimen Preparation." CSF specimens showing turbidity should be centrifuged for 10 min at 1400 x g after heating, and prior to testing. The supernatant fluid is to be used as the test specimen. Urine specimens (see "Specimen Preparation") that yield uninterpretable or suspected false positive results may be filtered using a MILLIPOREIM Millex-HA 0.45 filter (#SLHA-0250S) and retested with the reactive test latex (es) and the corresponding control latex(es). Colony confirmation samples that yield uninterpretable or atypical agglutination reactions may require a 1:10 dilution of the supernatant with Control - reagent. Care should be taken to avoid bacterial suspensions greater than a McFarland #1 turbidity standard. Latex agglutination tests of urine have potential usefulness in providing supportive evidence for the presence of a pathogenic organisms in infants presently with fever with localizing signs.22 However, the excretion of urinary antigen may persist in infants for 3 - 14 days, or up to 30 days, following the administration of Haemophilia influenzae type b (Hib) vaccines. Frequency and duration of antigenuria depends on the patient and the type of vaccine given. Patient vaccination history is important for the correct interpretation of positive urinary Hib antigen results in vaccinated infants. Cross-reactions have been produced by some strains of the viridans streptococci group (i.e., S. mitis, S. sanguis II, and S. salivarius) with the Directigen S. pneumoniae latex reagent. Additional tests may be needed to differentiate the viridans streptococci group from S. pneumoniae. A Gram stain should be performed prior to use with the colony confirmation procedure. Certain bacterial isolates obtained from blood culture plates, which are morphologically similar to, but Gram stain inconsistent with the suspect colony, may exhibit non-specific agglutination and/or false positive reactions (i.e., beta-hemolytic nonpathogenic Neisseria species) with the group B Strep latex. REFERENCES Becton Dickinson and Company, Directigen Meningitis Combo Test, 2/2015 |  |  |  |