ETSU Physicians and Associates

# ETSUHealth

MEAC POL

| Policy Title: Acceava Strep A |  |
| --- | --- |
| Date Effective/Author: 11/30/2022 CEF |  |

| Reviewed by: | Date: | Revision Number: (i.e. A,B,C) |
| --- | --- | --- |

APPROVED BY:

Signature:	  ____________________________________________________

Administrative Director, ETSU Clinical Labs		Date

Medical Director, MEAC POL

# Signature:	_____________________________________________________

# Lab Manager, ETSU Clinical Labs			Date

# Purpose:

The Acceava® Strep A is a rapid test to qualitatively detect the presence of Strep A antigen in throat swab specimens, providing results within 5 minutes. The test utilizes antibodies specific for whole cell Lancefield Group A Streptococcus to selectively detect Strep A antigen in a throat swab specimen to aid in the diagnosis of Group A Streptococcal infection.

Principle:

The Acceava® Strep A test is a qualitative, lateral flow immunoassay for the detection of Strep A carbohydrate antigen in a throat swab. In this test, antibody specific to Strep A carbohydrate antigen is coated on the test line region of the strip.  During testing, the extracted throat swab specimen reacts with an antibody to Strep A that is coated onto particles. The mixture migrates up the membrane to react with the antibody to Strep A on the membrane and generates a red line in the test region. The presence of this red line in the test region indicates a positive result, while its absence indicates a negative result. To serve as a procedural control, a red line will always appear in the control region if the test has been performed properly. If a red control line does not appear, the test result is not valid.

# Specimen Collection and Requirements:

Use of the sterile polyester swab supplied in the test kit is recommended.  Swabs from other manufacturers have not been validated.

Rayon transport swabs containing modified Stuart’s or Amies liquid medium can also be used with this product.

Swab posterior pharynx, tonsils and other inflamed areas.  Avoid touching the tongue, cheeks and teeth with the swab.

# Reagents & Equipment:

- Test Strips

- Sterile Polyester Swabs

- Disposable Extraction Test Tubes

- Reagent 1

- Reagent 2

- Positive Control

- Negative Control

- Timer

Storage and Stability

The kit can be stored at room temperature or refrigerated 36°-86°F / 2°-30°C. The test strip must remain in the closed canister until use. DO NOT FREEZE. The test strips (in their unopened canister) and the reagents are stable through the expiration date printed on the box.  Once the canister is opened, the remaining test strips are stable for 12 months.  Do not use beyond the expiration date.

###### Quality Control:

Internal Quality Control

Internal procedural controls are included in the test. A red line appearing in the control region (C) is an internal positive procedural control. It confirms sufficient specimen volume and correct procedural technique. A clear background is an internal negative background control. If the test is working properly, the background in the result area should be white to light pink and not interfere with the ability to read the test result.

External Quality Control

Quality control requirements must be performed in accordance with local, state, and federal regulations or accreditation requirements. External quality control must be performed every 30 days and with each new lot of test kits. External positive and negative controls are supplied in the kit. Alternatively, other Group A and non-Group A Streptococcus ATCC reference strains may be used as external controls. Some commercial controls may contain interfering preservatives; therefore, other commercial controls are not recommended.

Procedure for External Quality Control Testing

1. Add 4 full drops of Reagent 1 and 4 full drops of Reagent 2 into an extraction test tube. Tap the   bottom of the tube gently to mix the liquid.

2. Add 1 full drop of positive or negative control solution into the tube, holding the bottle upright.

3. Place a clean swab into the tube. Rotate the swab 10 times in the tube. Leave the swab in the tube for 1 minute. Then press the swab against the side of the tube and squeeze the bottom of the tube while removing the swab so that most of the liquid stays in the tube. Discard the swab.

4. Continue with Step 4 of Test Procedure.

# Testing Procedure / Operating instructions:

Allow the test strip, reagents, and/or controls to reach room temperature (59-86°F / 15-30°C) prior to testing.

1. Hold the Reagent 1 bottle upright and add 4 full drops (approximately 240 μL) to an extraction test tube. Reagent 1 is red in color. Hold the Reagent 2 bottle upright and add 4 full drops (approximately 160 μL) to the tube. Reagent 2 is colorless. The addition of Reagent 2 to Reagent 1 changes the color of the solution from red to very faint yellow. Tap the bottom of the tube gently to mix the liquid.

2. Immediately add the throat swab into the tube of very faint yellow solution.  Vigorously mix by rotating the swab forcefully against the side of the tube 10 times.  Leave the swab in the tube for 1 minute. Then press the swab against the side of the tube and squeeze the bottom of the tube while removing the swab so that most of the liquid stays in the tube. Discard swab.

3. Remove the test strip from the canister and use it as soon as possible.  Best results will be obtained if the test is performed immediately.  Immediately close the canister tightly after removing the test strip.  Record initial opening date on the canister.  Once opened, the remaining test strips are stable for 12 months.

4.  With arrows pointing down, place the test strip into the tube of solution and then start the timer. If the procedure is followed correctly, the liquid should be at or just below the maximum line (MAX) on the test strip.  See the illustration.

5. Leave the strip in the tube and read the result at 5 minutes. The result is invalid after 10 minutes.

# Reporting Results:

POSITIVE*: Two distinct red lines appear. One line should be in the control region (C) and another line should be in the test region (T). A positive result indicates that Strep A was detected in the sample.

*NOTE: The intensity of the red color in the test line region (T) will vary depending on the concentration of Strep A present in the sample. Therefore, any shade of red in the test region (T) should be considered positive.

NEGATIVE: One red line appears in the control region (C). No apparent red or pink line appears in the test region (T). A negative result indicates that Strep A is not present in the sample, or is present below the detectable level of the test. It is recommended that a negative patient sample be confirmed by culture to confirm the absence of Strep A infection.  If clinical symptoms are not consistent with results, obtain another sample for culture.

INVALID: Control line fails to appear. Insufficient sample volume or incorrect procedural techniques are the most likely reasons for control line failure. Review the procedure and repeat the test with a new test strip. If the problem persists, discontinue using the test kit immediately and contact Alere Technical Support at 877-441-7440, Option 2.

###### Limitations:

1. The Acceava® Strep A test is for in vitro diagnostic use only. The test should be used for the detection of Strep A antigen in throat swab specimens only. Neither the quantitative value nor the rate of increase in Strep A antigen concentration can be determined by this qualitative test.

2. This test will only indicate the presence of Strep A antigen in the specimen from both viable and non-viable Group A Streptococcus bacteria.

3. It is recommended that a negative result obtained from this kit be confirmed by culture. A negative result may be obtained if the concentration of the Strep A antigen present in the throat swab is not adequate or is below the detectable level of the test.

4. Excess blood or mucus on the swab specimen may interfere with test performance and may yield a false positive result. Avoid touching the tongue, cheeks, and teeth4 and any bleeding areas of the mouth with the swab when collecting specimens.

5. As with all diagnostic tests, all results must be interpreted together with other clinical information available to the physician.

6.The Acceava® Strep A Test has been categorized as CLIA Waived only for the application of qualitative detection of Group A Streptococcal Antigen from throat swabs.  The application for the confirmation of presumptive Group A Streptococcal colonies recovered from cultures is not Waived.

X

Warnings and Precautions:

WARNING: Reagent 1 is harmful if swallowed or absorbed through skin. May cause eye irritation.

CAUTION: Reagent 2 may cause skin, eye and respiratory tract irritation.

- The positive and negative controls contain sodium azide (NaN3) as a   preservative.

- Do not interchange Reagent or external Control solution bottle caps.

- Do not interchange or mix components from different kit lots.

- Humidity and temperature can adversely affect results. S

References:

1. Manual of Clinical Microbiology, 6th Edition, ASM Press, p. 299-307.

2. Webb, KH. Pediatrics (Feb 1998), 101: 2, 2.

3. Bisno AL, Gerber MA, Gwaltney JM, Kaplan EL, Schwartz RH. Clinical Infectious Diseases (1997), 25, 574-83.

4. Shea, Y.R., Specimen Collection and Transport, in Clinical Microbiology Procedures Handbook, 	Isenberg, H.D., American Society of Microbiology, Washington, D.C., 1.1.1-1.1.30, 1992.

5. Nussinovitch, M, Finkelstein Y, Amir J, Varsano, I. Clinical Pediatrics (June 1999), 357-360.

6. Woods WA, Carter CT, Stack M, Connors Jr AF, Schlager TA, Southern Medical Journal (May 1999), 491-492.