ETSU Physicians and Associates

# ETSU Clinical Lab Policy

| Policy Title: | ## Influenza A & B – Waived Laboratory |
| --- | --- |
| Policy Manual Section: | ETSU Laboratory |
| Policy Number: |  |
| Date Effective/Author: Board Approval Date: | Feb 2017 / AMD |

Review/Revision History:

| Reviewed by: | Date: | Revision Number: (i.e. A,B,C) |
| --- | --- | --- |
| Carrie Freeman BS MT (AMT) | 4/18/2019 |  |

APPROVED BY:

Signature:	  ____________________________________________________

Administrative Director, ETSU Clinical Labs		Date

# Signature:	_____________________________________________________

# Lab Manager, ETSU Clinical Labs			Date

# Principle / Purpose:

BinaxNOW Influenza A&B test kit is used for the qualitative detection of influenza A and B nucleoprotein antigens in nasal swab specimens. Its intended use is for rapid differentiation of influenza A and B in symptomatic patients.

Specimen Requirements:

- Double polyester dry swabs

- *do not use calcium alginate swabs*

- Test swabs within one hour of collection.

# Reagents & Equipment:

- BinaxNOW Influenza A & B test kit

- Test devices

- Transfer pipettes

- Positive/Negative control material

- Elution solution vials

- NP swabs

- Timer

# QC:

- An untested strip has a blue line at the “control” position.  If the test is preforming properly, the blue line will turn pink on the strip for the positive control.

- After the test device is inoculated, the background color in the window should change from light pink to white within 15 minutes. Clearing of the background color is the negative control.

** Patient testing is to be repeated if either control fails. If the QC continues to fail, call ETSU Clinical Lab personnel for further assistance.

External Quality Control is performed by ETSU Clinical Lab personnel with each new shipment or lot of test kits.

# Sample Preparation Procedure:

Nasal Swabs –.

- Twist off the cap of the pre-filled elution solution test vial.

- Place patient swab in test vial and rotate vigorously three times in the liquid. Avoid making excessive bubbles during this.

- Press swab against the side of the vial as you turn the swab, squeezing on the sides of the vial and pulling the patient swab up and out of the vial at the same time. This will allow the sample to be removed from the swab and remain in the liquid in the vial.

- Discard the swab in a biohazard container.

- Proceed to testing procedure.

# Testing Procedure / Operating instructions:

NOTE: The following procedure is for the BinaxNOW Influenza A & B kits. INVALID results may occur if too little sample is added to the test.

- Pull lab order for the influenza from the patient chart located in the LIS.  The LIS will print a bar coded label with a specimen number.  This number is to identify the patient and test ordered.

- Remove device from the pouch just prior to testing and lay flat on work bench.

- Place patient label (LIS generated) on the test kit for proper identification.

- Fill pipette by firmly squeezing the top bulb and pipette tip into sample.  Release bulb while tip is still in sample.  This will pull liquid into the pipette.  Make sure there are no air spaces in the lower part of the pipette.

- See arrow on test device to find White Sample Pad.  SLOWLY (drop by drop) add entire contents of pipette (100 ul) to the MIDDLE of this pad by squeezing the top bulb.

- Immediately peel off adhesive liner from the test device.  Close and securely seal the device. Read result in window in 15 minutes. Results read before or after 15 minutes may be inaccurate.

# Results Interpretation:

*Note: Do not read test results before or after 15 minutes as they may not be correct.

- NEGATIVE SAMPLE, the BLUE Control Line in the BOTTOM THIRD of the window turns a pink to purple color.  No other line appears.

- POSITIVE A SAMPLE, the BLUE Control Line turns a pink to purple color.  A SECOND pink to purple Sample Line appears above it in the MIDDLE THIRD of the window. Any shade of pink to purple Sample Line indicates a positive result.

- POSITIVE B SAMPLE, the BLUE Control Line turns a pink to purple color.  A SECOND pink to purple Sample Line appears above it in the TOP THIRD of the window.  Any shade of a pink to purple Sample Line indicates a positive result.

- POSITIVE A & B SAMPLE, the BLUE Control Line turns a pink to purple color, AND two pink to purple Sample Lines appear above it in the MIDDLE and TOP thirds of the window.  Any shade of pink to purple Sample Lines indicates positive results.

- INVALID if the Control Line remains BLUE or is not present at all, whether a Sample Line(s) is present or not. Repeat invalid tests.

*NOTE: A co-infection with influenza A & B is very rare and any patient test that is positive for A and B is invalid and the test is to be repeat with a recollected specimen or tested by another method.

*Expected value:  Negative for both A &B

# Reporting Results:

All results must be hand entered into the LIS and approved.  This will allow the results to transfer over into the electronic patient chart and electronic billing to occur.

###### Troubleshooting / Corrective Actions:

Refer to BinaxNOW product instructions insert, call ETSU Clinical Lab personnel for further assistance, or call Binax at 1-800-257-9525

Limitations:

As with all laboratory tests, definitive diagnostic or therapeutic decisions should not be based on any single result or method.

A negative test result does not exclude infection with influenza A and /or B.  Therefore, the results obtained with the BinaxNOW Influenza A & B Test should be used in conjunction with clinical findings to make an accurate diagnosis.

The BinaxNOW Influenza A & B Test detects both viable and non-viable influenza A and B.  Test performance depends on the virus (antigen) in the specimen and may or may not compare with cell culture results performed on the same specimen.

Inadequate specimen collection or improper sample handling/transport may yield a false-negative result.

Performance of the BinaxNow Influenza A & B Test has not been established for monitoring antiviral treatment of influenza

Use of visibly bloody samples is not recommended with the BinaxNOW Influenza A & B Test.

Warnings and Precautions:

Universal Precautions are to be followed while handling all blood and body fluids.  All specimens and containers are to be disposed of in the biohazard containers.  Wear gloves at all times and disinfect work area before and after each patient.

References:

BinaxNOW Influenza A & B kit product instructions insert 2005 Binax, Inc.  rev. 2  2/16/06  Inverness Medical professional diagnostics