KPNW Urine Pregnancy Limitations Fact Sheet:

- The Osom Urine hCG cartridge test is ONLY capable of detecting whole molecule (intact) hCG.  It cannot detect the presence of free hCG subunits.

- The presence of hCG β core fragment (hCGβcf), the predominant variant of hCG in urine, has been demonstrated to cause FALSE NEGATIVES  in qualitiative POC hCG devices.

- TEST RESULTS SHOULD BE CONFIRMED USING A QUANTITATIVE hCG ASSAY PRIOR TO THE PERFORMANCE OF ANY CRITICAL MEDICAL PROCEDURE.

- Dilute urine specimens may not have representative levels of hCG.  Specific gravity is not measured when UCG is tested as POCT.

- For early determination of pregnancy the first morning specimen of urine is recommended since it usually contains the highest concentration of hCG.