KPNW NOVA Xpress STATSTRIP: Training Checklist

- Review test system components, Meter, Strips and Quality Control Material.

- Review Open and discard dates & reagent labeling for test strips and QC unless manufacturer date is sooner. (90 days QC, 180 for Strips – 3months minus 3 days 6months minus 6 days)

- Review QC frequency (Monthly (30days), Documented on Manual QC Log)

  - If it is found that QC is not run after an emergency in which the meter was used explain what to do: run QC immediately

- Demonstrate how to perform a Glucose test

  - When running QC: Scroll to the proper QC level

  - Place the test strip into the meter. Note the screen change – flashing blood drop

  - Discuss how to hold the meter to sample QC or Patients. No liquid should get in the test strip port.

  - Hold the meter correctly when sampling to ensure continuous flow of blood – no double dip

  - You must manually document to electronic medical record.

  - Clean the meter exterior being careful not to get liquid in the strip port.

- Perform both Positive and Negative levels of QC and document on QC Log.

- Review how to document patient results in KPHC (HIS)