Policy: 408-I-003-10

| RESPONSIBLE DEPARTMENT: LABORATORY | POLICY #: 408-I-003-10 | PAGE: 1 of 4 |
| --- | --- | --- |
| SUBJECT: Order of Draw for Multiple Tube Collections | EFFECTIVE DATE: June 2008 |  |
| REVIEWED/ REVISED: October 2018 |  |  |
| DISTRIBUTION: PHLEBOTOMY |  |  |

## Purpose:

Quality control and accurate test results begin with patient preparation and sample collection.  Following the recommended order of draw for blood specimens is important for specimen integrity.

## Reagents / Equipment:

Evacuated Blood Collection Tubes

Procedure:

- Evacuated Tube Method

When drawing multiple vacutainer tubes during a single venipuncture the following order of draw should be followed:

Tube Type				Additive

1.	Blood Culture – Aerobic		Sterile Specimen

2.	Blood Culture – Anaerobic		Sterile Specimen

3.	Light Blue*				Sodium Citrate

4.	Red/SST/Gold				Clot Activator

5.	Green 					Heparin

6.	Lavender/pearl			EDTA

7.	Gray					Sodium Fluoride/Potassium Oxalate

8.	Other Additive Tubes

*NOTE:  When using a winged blood collection set for venipuncture and a coagulation tube is the first tube to be drawn, a discard tube should be drawn first.  The discard tube must be used to fill the blood collection tubing dead space and to assure maintenance of the proper anticoagulation/blood ratio.  The discard tube should be a non-additive or Light Blue coagulation tube, and it need not be completely filled.

B.   Microtainer Collection Method

If multiple specimens are to be collected by skin puncture (heel or fingerstick), anticoagulant tubes must be collected first to avoid microclots from forming because of a prolonged collection.

1.	Blood gases

2.	Slides/smears

3.	EDTA tubes

4.	Other additive tubes

5.	Serum tubes

C.  Errors

1.  Anticoagulant cross contamination

- EDTA, a calcium agent can yield false magnesium and calcium results if drawn

before a red top tube and can falsely elevate potassium results if drawn before a green top tube.

- Oxalate can interfere with cell membranes.

- Fluoride alters cell morphology if drawn before a lavender top tube.

- Interference with coagulation testing can occur when plastic serum tubes containing a clot activator are drawn before the coagulation tube.

2.  Tube inversion

- Blood collected in tubes with anticoagulants must be thoroughly mixed with the anticoagulant to prevent clotting.

- DO NOT SHAKE THE TUBE

- Mix by inverting the tube 5 times so the anticoagulant is thoroughly mixed with the blood.

- If mixing is not thorough, a partial clot may form and could render the specimen               unacceptable for use.

- Serum Separator tubes should also be inverted.  There is a clot activator present and         inversion expedites the clotting process.

Precautions:

HANDLING BLOOD SPECIMENS REQUIRES USE OF LAB COAT AND GLOVES, AND SAFETY SHIELD FACE PROTECTION IN DESIGNATED AREAS. WASH HANDS AFTER REMOVING GLOVES.

References:

CLSI Guidelines, “Procedures for the Collection of Diagnostic Blood Specimens by Venipuncture”; Approved Standard-Sixth Edition.  H3-A6, 2007.

CLSI.  Collection of Diagnostic Venous Blood Specimens. 7th ed. CLSI standard GP41.

Wayne, PA: Clinical and Laboratory Standards Institute; 2017.

| Approved By: |  |  |
| --- | --- | --- |
|  |  | Date |
|  | Phlebotomy Supervisor |  |
|  | Phlebotomy Supervisor |  |
|  |  | Date |
|  | Medical Director |  |

Adopted or Date initiated:  June 25, 2008

Revised and effective date:  June 2008, September 2010, December 2010,

December 23, 2014, October 2018

Reviewed:

| DATE | REVIEWER | DATE | REVIEWER |
| --- | --- | --- | --- |

| Revision History |  |  |  |
| --- | --- | --- | --- |
| Version | Summary of Changes | Author | Date |
| Version 4 | Policy on Syringe method removed/no longer performed by phlebotomy staff | FV | 12/11/14 |
| 5 | Slightly revise format, updated approval box | glik | 10/4/16 |
| 6 | Added header, added clot activator interference, updated EDTA tubes, added reference. | olearj | 10/09/18 |