Procedure 
Dignity Health Central Coast Service Area

SUBJECT: Performing the Flush System Process for the Cell Washer

ORIGIN: Transfusion Services

NUMBER: 7540.BB.CC.243

| Applies to: |  |  |
| --- | --- | --- |
| Santa Maria Campus, Marian Regional Medical Center | Arroyo Grande Campus, Marian Regional Medical Center | French Hospital Medical Center |
| St. John’s Pleasant Valley Hospital | St. John’s Regional Medical Center |  |

Purpose	To provide instructions to clean and disinfect the cell washer, as well as to remove blockages due to salt crystallization.

Policy

- The flush process using 10% bleach will be performed as part of the weekly maintenance.

- The flush process using only distilled water will be performed as part of daily maintenance and when the cell washer has been idle for > 4 hours.

Materials

| Supplies | Equipment |
| --- | --- |
| - 12x75 tubes - 500 mL of fresh 10% bleach solution - 1L distilled or deionized water - Blood bank saline | - Cell washer |

Procedure

Flushing the System with 10% Bleach: Follow the steps below to clean and disinfect the cell washer.

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | Load the rotor with tubes, leaving every other position in the rotor empty and close the lid. |  |
| 2 | Prepare 500 mL of fresh 10% bleach solution by combining 479 mL of distilled or deionized water and 21 mL of 6.0% bleach in the designated container. |  |
| 3 | Prepare a container of 1 L of distilled or deionized water. |  |
| 4 | Disconnect the saline tube from the saline cube and place tubing in a container containing the prepared 10% bleach solution. |  |

| 5 | Press the Menu button on the Start screen. Select and run the “flush 1” program. Wait 5 minutes. |  |
| --- | --- | --- |
| 6 | Remove the saline tubing from the bleach solution and place in a container containing 1 L of distilled or deionized water. |  |
| 7 | Press the Menu button on the Start screen. Select and run the “flush 1” program. |  |
| 8 | Open the lid and dry the centrifuge area. |  |
| 9 | Reattach the saline tubing to the saline cube. |  |
| 10 | Press the Menu button on the Start screen. Select and run the “flush 1” program. |  |
| 11 | Open the lid and dry the centrifuge area. |  |
| 12 | Document the completed maintenance task on the Helmer Automatic Cell Washer Maintenance Record. |  |

Flushing the System with Distilled or Deionized Water: Follow the step below to flush the system with distilled or deionized water.

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | Load the rotor with tubes, leaving every other position on the rotor empty. Close the lid. |  |
| 2 | Disconnect the saline tube from the saline cube and place tubing in a container with 1 L of distilled or deionized water. |  |
| 3 | Press the Menu button on the Start screen. Select and run the “flush 1” program. Wait 5 minutes. |  |
| 4 | Remove the saline tubing from the distilled or deionized water container and reattach the saline tubing to the saline cube. |  |
| 5 | Press the Menu button on the Start screen. Select and run the “flush 1” program. |  |
| 6 | Open the lid and dry the centrifuge area. |  |
| 7 | Document the completed maintenance task on the Helmer Automatic Cell Washer Maintenance Record. |  |

References

- Helmer Scientific Automatic Cell Washing System Services Manual Ultra CWII

Associated Documents

- Helmer Automatic Cell Washer Maintenance Record