Procedure 
Dignity Health Central Coast Service Area

SUBJECT: Operating the Cell Washer

ORIGIN: Transfusion Services

NUMBER: 7540.BB.CC.241

| Applies to: |  |  |
| --- | --- | --- |
| Santa Maria Campus, Marian Regional Medical Center | Arroyo Grande Campus, Marian Regional Medical Center | French Hospital Medical Center |
| St. John’s Pleasant Valley Hospital | St. John’s Regional Medical Center |  |

Purpose	To provide instructions on how to properly operate the Helmer Ultra CWII cell washer.

Policy

- The interior of the cell washer must be wiped down after each use.

- If the cell washer has been idle for longer than 4 hours, the cell washer must be flushed with distilled or deionized water. Refer to “Performing the Flush System Process for the Cell Washer” procedure.

Materials

| Supplies | Equipment |
| --- | --- |
| - Blood bank saline - 12 x 75 tubes | - Cell washer |

Procedure

Operating the Cell Washer: Follow the steps below to operate the cell washer.

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | Load tubes evenly around the rotor and close the lid. |  |
| 2 | Select the Menu button from the Start screen. |  |
| 3 | Select the appropriate wash program. If…; Then select the program …. 3 wash cycles are needed; 12x75 35 sec 3x washes. 4 wash cycles are needed; 12x75 35 sec 4x washes. |  |
| 4 | Select the green checkmark to load the program. The Start screen appears with the loaded program name at the top of the screen. |  |
| 5 | Press the Start Program button to begin the wash program. |  |

| 6 | Press the Open Lid button once the cycle is complete. The “Program Successful” dialog box will display along with an audible beep when the wash program is complete. |  |
| --- | --- | --- |
| 7 | Wipe the interior of the cell washer down with dry paper towels after each wash cycle. |  |

References

- Helmer Scientific Automatic Cell Washing System Services Manual Ultra CWII

Associated Documents

- Performing the Flush System Process for the Cell Washer