Procedure 
Dignity Health Central Coast Service Area

SUBJECT: Measuring and Calibrating the Cell Washer Saline Volume

ORIGIN: Transfusion Services

NUMBER: 7540.BB.CC.244

| Applies to: |  |  |
| --- | --- | --- |
| Santa Maria Campus, Marian Regional Medical Center | Arroyo Grande Campus, Marian Regional Medical Center | French Hospital Medical Center |
| St. John’s Pleasant Valley Hospital | St. John’s Regional Medical Center |  |

Purpose	To verify the volume of dispensed saline and calibrate if needed as part of monthly preventative maintenance.

Materials

| Supplies | Equipment |
| --- | --- |
| - Blood bank saline - Graduated cylinder | - Cell washer |

Procedure

Measuring the Saline Volume: Follow the steps below to measure the saline volume.

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | Touch the login button on the Start screen. |  |
| 2 | Enter the Service User password: 46060. Press the green check mark when finished. |  |
| 3 | Select the System Settings button then select the tools button. |  |
| 4 | Select Ruler icon form the Service Menu screen. |  |
| 5 | Press the “36 mL” button. |  |
| 6 | Hold the graduated cylinder under the nozzle in the lid of the cell washer and press the green Start button. |  |

| 7 | Measure the liquid collected when the liquid has stopped flowing from the spout. Ensure the collected volume is 36 mL ± 1.8 mL. If the saline volume measured is…; Then…. Within acceptable range; - Cell washer can be used.. NOT within range; - Use the “+”, “++”, “-“ and “- -“ buttons to adjust the number to match the measured volume. |  |
| --- | --- | --- |
| 8 | Select green checkmark when calibration is complete. Use red arrow to return to the home screen. |  |
| 9 | Document the completed maintenance task on the Helmer Automatic Cell Washer Maintenance Record. |  |

References

- Helmer Scientific Automatic Cell Washing System Services Manual Ultra CWII

Associated Documents

- Helmer Automatic Cell Washer Maintenance Record