Procedure 
Dignity Health Central Coast Service Area

SUBJECT: Cleaning the Cell Washer and Cell Washer Rotor

ORIGIN: Transfusion Services

NUMBER: 7540.BB.CC.242

| Applies to: |  |  |
| --- | --- | --- |
| Santa Maria Campus, Marian Regional Medical Center | Arroyo Grande Campus, Marian Regional Medical Center | French Hospital Medical Center |
| St. John’s Pleasant Valley Hospital | St. John’s Regional Medical Center |  |

Purpose	To provide instructions for cleaning the cell washer and the cell washer rotor.

Policy	The cell washer and cell washer rotor will be cleaned as needed and as part of the following maintenance schedule:

- Cell washer interior: daily maintenance

- Cell washer exterior: monthly maintenance

- Cell washer rotor: weekly maintenance

Materials

| Supplies | Equipment |
| --- | --- |
| - Dish soap - Paper towel | - Helmer Ultra CWII Cell Washer |

Maintenance	Maintenance will be performed following the daily, monthly, and weekly maintenance schedule.  Completed maintenance tasks will be documented on the Helmer Automatic Cell Washer Maintenance Record.

Procedure

Cleaning the Cell Washer: Follow the steps below to clean the exterior and interior of the cell washer

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | Prepare a soapy water solution by combining tap water with a few drops of dish soap. |  |
| 2 | Clean the exterior of the cell washer with soapy water. |  |
| 3 | Remove soap with a damp paper towel and towel dry. |  |
| 4 | Remove rotor and splash guard cap. Inspect the interior of the cell washer and ensure the bowl is free and clear of debris. |  |

| 5 | Wipe the centrifuge bowl and splash guard ca with a damp paper towel. Place the splash guard cap back in the centrifuge. |  |
| --- | --- | --- |
| 6 | Dry the centrifuge bowl and splash guard cap with a dry paper towel. Place the splash guard cap back in the centrifuge. |  |
| 7 | Document the completed maintenance task on the Helmer Automatic Cell Washer Maintenance Record. |  |

Cleaning the Cell Washer Rotor: Follow the steps below to clean the cell washer rotor.

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | Remove rotor from cell washer. |  |
| 2 | Run warm tap water directly into the top of the rotor for several minutes. Ensure that water is flowing freely out of all the ports. If water is …; Then …. Flowing freely; - Proceed to step 3. Not flowing freely; - Gently slide the lid release tool into the fill port from the outside towards the center of the rotor and slide the lid release tool in and out several times. - Repeat step 2 until water is flowing freely.. |  |
| 3 | Dry the rotor with a paper towel and return to the cell washer. |  |
| 4 | Document the completed maintenance task on the Helmer Automatic Cell Washer Maintenance Record. |  |

References

- Helmer Scientific Automatic Cell Washing System Services Manual Ultra CWII

Associated Documents

- Helmer Automatic Cell Washer Maintenance Record