UW Medicine - Pathology

6000-02-02-09

IHC Microtomy Procedure

Adopted Date: 12/03/08 Review Date: 12/14/10 Revision Date: 06/13/11

PURPOSE

Instructions for properly cutting tissue sections for the Immunohistochemistry laboratory.

SCOPE

Histology Department

PROCEDURE

- Use either tap water or deionized water for the water bath.

- Use Positive charged slides and cut sections at 4 microns.

- Select appropriate pre-cut control slide or control block, whichever is applicable.  Place in proper order in slide printer.

- Place patient sections on the bottom half of the slide under the control, to ensure that the whole section is covered by the reagents during staining.

Note:  Be sure tissue sections are placed in appropriate "window" for Bond instrumentation.

- When possible, maintain the same tissue orientation on each slide.

- On requests that contain both antibody and unstained orders, cut antibody slides first and then the remaining unstained slides, numbering the slides if there is scant tissue.  Then cut the H&E recut last.  Label unstained with the case number only.

- Dry unstained slides for 30 minutes only in the 60oC drying oven in a metal slide rack, and stain the H&E recut slide on the stainer.

Note:  Unstained slides with no orders are filed "unbaked".  Give them to IHC for filing.

- Mark the order as complete in PowerPath after printing the H&E recut label.

Written By:					Director Approval:

_______________________		___________________________

Histology Supervisor