UW Medicine - Pathology

6000-02-07-02

Cryostat Maintenance and Decontamination Procedure

Adopted Date: 04/06/11 Review Date: Revision Date: 05/03/11

PURPOSE

To disinfect the cryostat on a regular basis.

SCOPE

This procedure applies to all staff performing frozen sectioning.

PROCEDURE

Principle:

According to CAP regulations, a cryostat should be defrosted and decontaminated with a tuberculocidal disinfectant at a time interval appropriate for the institution.  The regulations also state that the cryostat must be clearly marked as contaminated if a frozen section is performed on tissue from a patient known or suspected to be positive for HIV, hepatitis B or C, SARS-related coronavirus, prion disease such as Creutzfeldt-Jakob disease, or mycobacterial or systemic fungal disease.  It must then be decontaminated before further use.

General Precautions:

- Fluid resistant barrier gowns or lab coats, puncture and penetration resistant gloves and eye protection must be worn when performing cryostat disinfection procedures.  Virugard can be corrosive and can cause skin damage.

- Prior to manipulating the knife and specimen, or changing the specimen or knife, and during breaks, always lock the hand-wheel and cover the cutting edge with the tissue roll plate.

Reagents:

- ViraGuard (Veridian Corporation)

- 70% Ethanol

- 100% Ethanol

Procedure:

  - Daily Cryostat Cleaning:

    - Verify that no patient samples are present in the chamber.

    - Remove used blades / knives from their holder.  Do not clean with blade in place. Although not a requirement, steel mesh gloves should be worn when changing knife blades.

    - Dispose of blades in a provided sharps container.

    - Wipe down all utensils (pencils, forceps, brushes, gauze, etc) with 70% alcohol.  Set aside, returning to the chamber after cleaning.

    - Remove all debris from the chamber.  Debris must be removed because organic materials (blood and proteins) may contain high concentrations of microorganisms and could possibly inactivate the chemical disinfectant or prevent access to contaminated surfaces.  It should be treated as biohazardous waste and disposed of according to the policies and procedures of the institution.

    - 70% ethyl alcohol poured on a cloth should be used to clean the cryostat and provide some disinfection capabilities.  Be sure to treat all surfaces that might touch fresh tissue including the cryostat lid.

  - Chemical Disinfection / Decontamination:

      - Disinfection / Decontamination should be performed at an interval deemed appropriate with the usage of the cryostat.

        - If the Cryostat is used daily, it must be decontaminated weekly.

          - Pour disinfectant (Viraguard) onto surfaces or absorbent disposable towels and wipe the contaminated surfaces and leave wet for 10 minutes.  When disinfecting the Cryostat using Viraguard, it is not necessary to thaw the Cryostat.

          - Accessories with multiple parts, such as the disposable blade holder, knife holder and their respective bases, must be taken apart, disinfected with Viraguard and dried thoroughly.

          - Do not create aerosols by spraying disinfectant (or anything else) in an open cryostat chamber.

          - Any disposable material used in the disinfection process must be disposed of in accordance with the Compromised Specimen procedure, 6000-01-04-01, and Safety, Infection Control Management policy, 100-01-04-01.

  - Following Disinfection:

            - After the disinfection procedure is complete, absolute ethyl alcohol can be used to remove excess moisture from surfaces.

            - The liquid waste container should be emptied in accordance with the polices and procedures of the institution.  Before replacing the instrument, add a small amount of liquid bleach to the empty container.

  - Maintenance:

            - Lubricate the main cylinder after decontamination.

            - Move the chuck assemble to its further forward position and apply a drop of lubricant oil.  Move the assembly back to the home position.

              - Notes:

- The germicidal activity of ethyl alcohol is most effective in the 70% range because it can penetrate tubercle bacteria and it has an advantage over isopropyl alcohol of being able to kill hydrophilic viruses.

Note: The cryostat must be defrosted and decontaminated with a tuberculocidal disinfectant at an interval appropriate for the institution; this must be a weekly for instruments used daily.  Trimming and sections of tissue that accumulate inside the cryostat must be removed during decontamination.  Although not a requirement, steel mesh gloves should be worn when changing knife blades.

REFERENCE

Carson, Freda, Histotechnology a Self-Instructional Text 2nd edition, ASCP Press.  Chicago, IL ASCP Press, 1990.

ViraGuard MSDS

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________