UW Medicine - Pathology

6000-02-07-07

Prisma H&E Stainer Reagent QC / Maintenance Procedure

Adopted Date: 12/20/10 Review Date: Revision Date:

PURPOSE

To provide direction for changing solutions on the Prisma H&E stainer.

SCOPE

Histology staining

PROCEDURE

Tissue Tek Prisma Stainer Changing Protocol

Daily After Workload Has Been Stained:

- Dump 100% #16 into waste container, clean, fill with fresh Flex 100.  Replace at location #18.

- Move 100% #17 to 100% #16 location.  Move 100% #18 to 100% #17 location.

- Run a control slide with current date for QC.

Biweekly After Workload Has Been Stained:

  - Clean Load bins and Water stations as needed.

  - Dump Xylene #1 into waste container, clean, fill with fresh Xylene.  Replace at Xylene #3 location.

  - Move Xylene #2 to #1 location and Xylene #3 to #1 location.

  - Dump 100% #4 into waste container, clean, fill with fresh Flex 100 and replace at 100% #5 locaton.

  - Move 100% #5 to the #4 location.

  - Dump HALF of the Hematoxylin #8 into "Combined Dyes" container, fill with fresh Hematoxylin (#7211) and replace.

  - Dump Clarifier #10 down drain (with water running), clean, fill and replace with fresh Clarifier 2.

  - Dump HALF of the Eosin #15 into Eosin waste container, fill with fresh Eosin-Y and replace.

  - Dump 100% #16 into waste container, clean, fill with fresh Flex 100 and replace to 100% #18 location.

  - Move 100% #17 to 10% #16 location.  Move 100% #18 to 100% #17 location.

  - Dump Xylene #19 into waste container, clean, fill with fresh Xylene and replace to Xylene #21 location.

  - Move Xylene #20 to Xylene #19 location.  Move Xylene #21 to Xylene #20 location.

  - Top off or replace any other containers as needed.

  - Run a control slide with current date for QC.

Weekly After Workload Has Been Stained:

- Follow steps for biweekly change, but COMPLETELY dump Hemotoxylin and Eosin, fill with fresh Hematoxylin (#7211) and Eosin-Y and replace containers with fresh containers.

- Dump 95% #6 and 95% #14 into waste container, clean, fill with fresh Flex 95 and replace.

- Dump Bluing #12 down drain (with water running), clean, fill with fresh Bluing and replace.

- Top off or replace any other containers as needed.

- Run a control slide with current date for QC.

Written By:					Director Approval:

_______________________		___________________________

Histology Supervisor