UW Medicine - Pathology

6000-02-07-11

OHAUS Scale Usage and Maintenance Procedure

Adopted Date:4/25/2011 Review Date: Revision Date:

PURPOSE

To ensure the proper usage and maintenance of the OHAUS scale.

PROCEDURE

Materials:

- Scale

- Stainless steel pan

- Battery or AC Power Connection

- Weigh boats

- Chemicals/powders

- Spatula

- Traceable Certified Weight set (Troemner)

General Use:

- Press the "On/Off Zero" button briefly to turn scale on.

- When the scale is turned on, all display segments will be displayed for a few seconds, indicating that the unit is self adjusting to zero.  Once the "0" is displayed, the scale is ready for use.

- To tare the scale, place measuring boat on the stainless steel surface and press the "On/Off Zero" button to zero the scale again.

- Press the "Unit/Cal" button briefly to change the weighing unit (gram, oz, lb or N).

- Add desired chemical/powder to weigh boat until desired unit of weight is displayed.

- Remove the boat, clean stainless steel surface and turn unit off by pressing and holding the "On/Off Zero" button.

Calibration Maintenance:

(for best results calibrate the scale at regular intervals.  This is especially important if the scale is in use for prolonged periods).

- Press and hold the "Unit/Cal" button to begin the calibration process. The display will show CAL

- The calibration process can be aborted by turning the scale off.

- Press the "On/Off Zero" button to capture 0. The display shows -C- while the scale stores the zero load signal.

- The display will show C xxx where xxx is the calibration weight in grams.

- Place the appropriate calibration weight on the platform.

- Press the "On/Off Zero" button.

- The display shows -C- while the scale stores the calibration point signal.

- After calibration, the display returns to the normal weighing mode.

- The message CAL E will appear if the calibration steps are not followed or the wrong weight was used.

RELATED DOCUMENTS:

Histology Equipment Maintenance Log

REFERENCE:

OHAUS Compact Series Instruction manual

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________