Date:	_3/23/23

Affected SOPs: _ Vision Daily Maintenance and Vision Monitoring Liquid Levels.

Description:

As part of the second part of the implementation of the Vision Max, the Riley Blood Bank Vision Max and original Vision will be converted from an external/internal waste container to a waste line to a waste outlet in the floor.

Summary of Changes:

On 3/23/23, the Max 2 at RHBB had the waste changed to external waste line from a waste management. This change will be a deviation from daily maintenance on the instrument to not empty the waste, but to only refill the saline and DI water.

The SOP will need to be updated to include the use of the internal/external waste containers as well as the waste being sent to the drain.

Only RHBB has a drain for the Vision, all other instruments will have containers for waste.

Implementation of Change

This change was reviewed with the team via MTS training document.

The team will be notified via email of the change.

Both SOPs and the form revision have been drafted for implementation no later than June 2023.

Jayanna Slayten	3/23/23

BB Supervisor						Approved By:  BB Medical Director

Jayanna Slayten, MS, MT(ASCP)SBBcm		Diganotes approval Dr. LePage, 3/24/23@1700