## Slide 1: Precision Xceed Pro
        Quality Control Requirements
        Blood Glucose and Ketone Measurements
        Limitations of Precision Xceed Pro
        Docking and Recording Results.

## Slide 2: Clinical Significance
Point of care blood glucose monitoring systems provide rapid results
required to make therapeutic decisions.
This type of testing; however, supplements, rather than substitutes
for testing in the clinical laboratory.

## Slide 3: Basic Operation
Press On/Off to turn on monitor.
Press 1 to select Patient Test.(Press 2 to select Quality Control.)
Manually enter operator ID.
Scan Patient Visit ID on patient wristband. Clinic patient ID is chart number. Do not enter manually.
  (Scan QC barcode on bottle).
Scan strip barcode.
Insert strip until “Strip Inserted” appears on monitor. Use each test strip only once.
Apply drop of blood directly from patients’ finger, syringe or 1 drop of QC
to target area on strip.
Monitor will beep and display “Sample Accepted”
Do not move instrument until result appears and strip has been removed.
Clean Instrument after use, and between patients.

## Slide 4: Use Alcohol Wipe to Clean Instrument after Each Use

## Slide 5: Quality Control

## Slide 6: Quality Control-Testing Process
Use low and high control solution.
  Record open date and expiration date on bottle. Expires 90 days after opened.
  Gently mix controls 3-4 times.
  Expel a drop, discard and wipe tip before placing drop on strip.
  Instrument should remain flat to avoid liquid in port.
  Do not pick up instrument until strip had been removed.
  Follow same procedure for Glucose and Ketone QC.

## Slide 7: Quality Control Troubleshooting
Mix the controls
3-4 times before testing.

## Slide 8: Instrument Ranges
Glucose
  20-400 mg/dL
Ketone
  0.0-8.0 mg/dL

## Slide 9: Patient Testing
Verify Physician Order
Verify Patient Identification
Clean the site
Perform Finger stick
Wipe away the first drop.
Apply blood
Test strip is designed to allow top-fill or end-fill.
Test will not start until sufficient blood is applied.
There will be an audible beep, and analysis will begin.
Note: If you hold the Scan button for three seconds, the scanner stops. Reposition scanner and try again.

## Slide 10: Action Ranges for Blood Glucose
Expected glucose ranges for a non-diabetic, non-pregnant fasting adult
is under 100 mg/dL.
Action Range: If a glucose is below 60 mg/dl or above 200 mg/dL and not covered
by current physicians orders, the physician will be notified.
CRITICAL VALUES: <40 or>400
<40 repeat test to verify.
>400 do not repeat
****Must be verified with a LABORATORY confirmation.
Order a STAT glucose and call the lab.*****

## Slide 11: HYPERGLYCEMIC PATIENTS
For patient glucose values >400mg/dL follow these steps.
  Order Laboratory Glucose STAT.
  Call Laboratory .
  Do not use glucometer until LABORATORY glucose result is less than 400mg/dL.
  If DKA Protocol has been initiated, notify lab.
  The glucometer will not give accurate results. This could result in harm to the patient or even cause death.

## Slide 12: Critically Ill Patients
  The Precision Xceed Pro cannot be used to monitor critically ill patients.
Critically ill patients at this facility are defined as: hyperglycemic- hyperosmolar state (with or without ketosis), severe dehydration, hypotension or Shock.
CBGs (capillary blood glucose) POC test cannot be used until Lab result is 400 or less.

## Slide 13: Notes for Ketones
Normally ketone levels are expected to be less than 0.6mmol/L.
Ketone results between 0.6 and 1.5mmol/L and the blood glucose is 300mg/dL or higher, could indicate development of a problem requiring medical intervention.
Ketone results higher than 1.5mmol/L and glucose result is 300mg/dL or higher, this may indicate the risk of developing diabetic ketoacidosis (DKA).

## Slide 14: Replacing the Batteries
A Low Battery can cause erroneous results. Change batteries when this warning appears on the monitor.
Batteries may be obtained from materials department
Turn the Xceed Pro Monitor off by pressing the On/Off button on the keypad.
Press the tab on the battery compartment slide cover down to remove.
Remove batteries.
Insert new batteries and replace cover.

## Slide 15: Docking instrument after use automatically transfers the result to patient electronic record, eliminating any errors in reporting. Clinic personnel will manually enter results to EMR.
Docking station uploads patient result to EMR. The docking station does not charge the batteries.

## Slide 16: Call the Laboratory ext 1428