**Hopi Health Care Center**

**Polacca, Arizona**

**Laboratory Department**

**Policy & Procedure**

|  |  |  |  |
| --- | --- | --- | --- |
| **Title: Urine Pregnancy Testing (hCG)** | | | |
| **Responsible Person: Kendrick Fritz, Laboratory Supervisor** | | | |
| **Standards/Regulations: WT.01.01.01, WT.04.01.01, and WT.05.01.01** | | | |
| **Distribution: Laboratory,** **ED/UC, OPD** | | | |
| **Original Effective Date:**  **05/2016** | **Reviewed/Revised:** | **Review Interval:**  Annual | **Due for Review:**  05/2017 |
| **HHCC Website: Chapter: Departmental Policies and Procedures**  **Folder: Laboratory Policy & Procedure** | | | |

**POLICY**

1. Urine pregnancy test will be done utilizing the Consult Diagnostics hCG Urine Cassette test kit.
2. Only nursing and medical staff who have had documented and demonstrated competency assessment for the point of care testing locations will perform urine pregnancy testing.
3. Urine pregnancy testing is for screening purposes only.

**INTENDED USE**

The Consult Diagnostics hCG Urine Cassette is a rapid chromatographic immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in urine as an aid in the early detection of pregnancy.

**SUMMARY**

Human chorionic gonadotropin (hCG) is a glycoprotein hormone produced by the developing placenta shortly after fertilization, In normal pregnancy, hCG levels continue to rise as early as 7 to 10 days after conception. hCG levels continue to rise rapidly frequently exceeding 100 mIU/mL by the first missed menstrual period., and peaking in the 100,000-200,000 mIU/mL range about 10 – 12 weeks into pregnancy. The appearance of hCG in urine soon after conception, and its subsequent rapid rise in concentration during early gestational growth, make it an excellent marker for early detection of pregnancy.

The Consult Diagnostics hCG Urine Cassette is a rapid test that qualitatively detects the presence of hCG at the sensitivity of 20 mIU/mL in urine. The test utilizes a combination of monoclonal and polyclonal antibodies to selectively detect elevated levels of hCG in urine. At the level of claimed sensitivity, the Consult Diagnostics hCG Urine Cassette shows no cross-reactivity interference from the structurally related glycoprotein hormone hFSH, hLH and hTSH at a high physiological levels.

**PRINCIPLE**

The test utilizes a combination of antibodies including mouse monoclonal anti-hCG antibodies and goat polyclonal anti-hCG antibodies to selectively detect elevated levels of hCG. The assay is conducted by adding urine specimen to the specimen well of the test cassette and observing the formation of colored lines. The specimen migrates via capillary action along the membrane to react with the colored conjugate.

Positive specimens react with the specific colored antibody conjugates and form a colored line at the test line region of the membrane. Absence of this colored line suggests a negative result. To serve as a procedural control, a colored line will always appear at the control line region if the test has been performed properly.

**MATERIALS REQUIRED**

1. Consult Diagnostics hCG Urine Cassette – test cassette contains anti-hCG gold conjuagtes and anti-hCG coated on the membrane.
2. Disposable pipette (contained in kit)
3. Clock or Timer
4. Collection Cup
5. Positive and Negative Controls
6. Procedure card

**REAGENT STABILITY & STORAGE**

1. Store Consult Diagnostics hCG Urine Cassette at room temperature (2°-30°C) in the sealed pouch.
2. Test cassettes are stable until the expiration date printed on the sealed pouch.
3. Do not use beyond expiration date.
4. DO NOT freeze.

**SPECIMEN COLLECTION**

1. Urine specimen collected in a clean, dry, sterile container without preservatives.
2. The first morning specimen of urine is recommended since it usually contains the highest concentration of hCG. However, specimens collected at random may also be used.
3. Urine specimen may be stored at room temperature (15°-30°C) for up to 2 hours, or refrigerated at 2°-8°C for up to 48 hours.
4. If specimen has been stored refrigerated, allow it to warm to room temperature before use.

**QUALITY CONTROL**

Internal Quality Control

The appearance of the control band in the results window is an internal positive procedural control which validates the following:

1. *Test System*: The appearance of the red control band assures that the detection component of both the test line and control line is intact, that adequate sample volume was added and that adequate capillary migration of the sample has occurred. Also, verifies proper assembly of the Test Device.
2. *Operator*: The appearance of the control band indicates that an adequate volume of fluid was added to the sample well for capillary migration to occur. If the control band does not appear at the read time, the test is invalid.
3. The clearing of the background in the results area may be documented as a negative procedural control. It also serves as an additional capillary flow control. At the read time, the background should appear white to light gray and not interfere with the reading of the test. The test is invalid if the background fails to clear and obscures the observation of a distinct control band.
4. Results of the internal quality control will be recorded with each patient test.

External Quality Control

1. Two levels of quality control need to be tested, negative and positive.
   1. KOVA-Trol I – Negative Control
   2. KOVA-Trol III – Positive Control.
2. The frequency of external negative and positive controls to be tested are:
   1. Each new lot number,
   2. Each new shipment,
   3. Each new untrained operator,
   4. And as required for:

i. Laboratory - weekly on Sunday and for (a) through (c) from above.

ii. Nursing Point-of-Care Testing – every two weeks (bimonthly), with each new lot number of test kits and each new shipment..

1. The following control results must be obtained before reporting patient results:
   1. Level 1 Negative Control: No red band in the test window (T) and one red band in the control window (C).
   2. Level 2 Positive Control: One red band in the test window (T) and one red band line in the control window (C).
   3. Internal Positive Control:A red band in the control window (C) confirms that a sufficient sample volume was delivered and that the correct procedure was used.
   4. Internal Negative Control: A white to light gray background area indicates that the test is working properly.
2. Results of the quality control must be recorded in the QC logbook.

**PROCEDURE – See the “Procedural Card” at end of procedure.**

**Allow the test cassette, urine, and/or controls to equilibrate to room temperature 15-30oC prior to testing.**

1. Remove the test cassette from the sealed pouch and the pipette from the pouch. Use as soon as possible.
2. Place the device on a flat surface. Squeeze the bulb of the pipette and insert the barrel into the patient sample. Release the bulb and draw up enough sample to dispense 3 drops of urine.
3. Hold pipette vertically and transfer 3 full drops of urine (approx.. 100 uL) to the specimen well (S) of the test cassette, and then start timer. Avoid trapping air bubbles in the specimen well.
4. Discard the pipette in a biohazardous waste container.
5. Read results at 3 minutes.

Note: Results should be read between 3 to 5 minutes only. A result seen after these times could be indicative of a low hCG level and recommended that the test be repeated with a new sample 48-72 hours or that an alternate confirmation method is used.

**INTERPRETATION OF RESULTS**

1. Positive Result: Two separate red lines, one at (T) – Test area, and the other at the (C) – Control area, are visible in the results window, indicating that the specimen contains detectable levels of hCG. While the intensity of the test band may vary with different specimens, the appearance of 2 distinct bands should be interpreted as a positive result.
   1. Note: If the test band appears very faint, it is recommended that a new sample be collected 48 hours later and tested using another Consult Diagnostics hCG Urine Cassette or that an alternate confirmation method is used..
2. Negative Result: If no appearant red or pink line appears at the (T) – Test area and a red band is visible at the (C) - Control area, the test is considered negative, indicating that the specimen contains no detectable level of hCG.
3. Invalid Result: If no band appears at the (C) - Control region, the test is invalid. The test is also invalid if incomplete or beaded bands appear at the (T) Test or (C) Control region. The test should be repeated using another device.
   1. Note: The Insufficient specimen volume or incorrect procedural technigues are the most likely reason for control line failure. Review the procedure and repeat the test with a new test cassette. If problem persists, discontinue the test kit immediately and contact Technical Support at (866) 216-0094.

**RESULTS REPORTING**

Reference Range: Negative results are expected in healthy non-pregnant women and health men.

Note: The Consult Diagnostics hCG Urine Cassette has a sensitivity of 20 mIU/mL in urine and is capable of detecting preganacy as early as 1 day after the first missed menses.

**DOCUMENTATION OF PATIENT AND QUALITY CONTROL RESULTS**

Results must be documented along with the initials of personnel performing the test and the date the test was performed. A functional audit trail must be maintained that allows retrieval of results.

1. Results are to be recorded on the test log.
2. Document date test was performed, sign or symptom, provider, and two patient identifiers.
3. The initials of point of care testing personnel performing patient testing must be documented on log.
4. Results must be entered into E.H.R. using the POC Lab Entry button. See the “Electronic Health Record POC Lab Entry Button for Entering Point of Care Test Results Procedure” for detailed instructions.

**LIMITATIONS**

1. Very dilute urine specimens, as indicated by a low specific gravity, may not contain representative levels of hCG. If pregnancy is still suspected, a first morning urine specimen should be collected 48 hours later and tested
2. False negative results may occur when levels of hCG are below the sensitivity level of the test. When pregnancy is still suspected, a first mornig urine specimen should be collected 48 hours later and tested.
3. Very low levels of hCG (less than 50 mIU/Ml) are present in urine shortly after implantation. However, because a high significant number of first trimester pregnancies terminate for natural reasons. Specimens tested as positive during the initial days after conception may later be negative due to natural termination of the pregnancy. Overall, natural termination occurs in 22% of clinical unrecognized pregnancies and 31% of other pregnancies. In the presence of weakly positive results, it is good laboratory practice to sample and test again after 48 hours.
4. A number of conditions other than pregnancy, including trophoblastic disease and certain non-trophoblastic neoplasms including testicular tumors, prostate cancer, breast cancer, and lung cancer, cause elevated levels of hCG. Therefore, the presence of hCG in urine specimen should not be used to diagnose pregnancy unless these conditions have been ruled out.
5. As with any assay employing mouse antibodies, the possibility exits for interference by human anti-mouse antibodies (HAMA) in the specimen. Specimens from patients who have received preparations of monoclonal antibodies for diagnosis or therapy may contain HAMA. Such specimens may cause false positive or false negative test results.
6. This test provides a presumptive diagnosis for pregnancy. A confirmed pregnancy diagnosis should only be made by a physician after all clinical and laboratory findings have been evaluated.

**REFERENCE**

*Consult* Diagnostics hCG Urine CassetteInsert, PSS World Medical, Inc.

**This Policy & Procedure was originated, reviewed and approved by the following:**

\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Kendrick Fritz, Supervisory Medical Technologist Date

\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Noelle E. Blue Arm, M.D., Medical Laboratory Director Date