Purpose:

This procedure describes how to replace old labels and ribbons on the Hematrax printer with new labels and ribbon

Procedure:

| Step |  | Action |  |
| --- | --- | --- | --- |
| Part A: Remove Label and Ribbon from Printer |  |  |  |
| 1 |  | - Turn Power switch off - Power switch is located back of printer |  |
| 2 |  | - Open Side Panel |  |
| 3 |  | - Release the printhead located at the front by pressing the latch at the lower right side |  |
| 4 |  | - Rewind label back through printer |  |
| 5 |  | - Remove the label spool from printer by flipping the latch down and sliding the roll out |  |
| 6 |  | - If removing ribbon also then go to next step - If not replacing ribbon then go to Part B Installing labels and ribbon |  |
| 7 |  | - Lower the printhead but do not clamp down |  |
| 8 |  | - Firmly grasp the ribbon with both hands and tear it |  |
| 9 |  | - Remove the unused spool of ribbon from the below |  |
| 10 |  | - Push the printhead down and secure |  |
| 11 |  | - Wind up excess ribbon by rotating the spool pressure knob clockwise |  |
| 12 |  | - Remove the used ribbon by grasping the roll, depress the black knobs and pull the roll sideways and outwards |  |
| 13 |  | - Discard the used ribbon and used label spool |  |
| Part B: Installing Labels and Ribbon in Printer |  |  |  |
| 1 |  |  | - Turn Power off |
| 2 |  |  | - Open side panel |
| 3 |  |  | - Open new roll of ribbon and labels |
| 4 |  |  | - Using a dust can spray the printhead, sensor area to remove any dust and - debris |
| Step |  |  | Action |
| Part B: Installing Labels and Ribbon in Printer continued. |  |  |  |
| 5 |  |  | - Slide in new roll of labels |
| 6 |  |  | - Release printhead - Note: Keep hands clear from touching the printhead |
| 7 |  |  | - Slide the label through the black guide - Note: Adjust label guide to proper width if need |
| 8 |  |  | - Feed2-3 labels through the front of printer |
| 9 |  |  | - Red dot shines from the sensor under the label |
| 10 |  |  | - Adjust sensor location to about 1 inch from label edge - Note: If red laser sensor is dirty use a alcohol wipe to remove any dust and debris |
| 11 |  |  | - Lower printhead but do not secure |
| 12 |  |  | - Prepare ribbon by unrolling part of the blue leader |
| 13 |  |  | Insert ribbon onto spindle - Note: Be sure to install ribbon all the way to spindle edge |
| 14 |  |  | - Pull blue ribbon leader under printhead |
| 15 |  |  | Separate 2 labels from spool - Remove label but save the backing |
| 16 |  |  | - With the shiny side of backing facing down, attach the blue ribbon leader |
| 17 |  |  | - Wrap label backing around the ribbon take up spool |
| 18 |  |  | - Wrap the ribbon around the spool 4-5times |
| 19 |  |  | - Push the printhead down and secure |
| 20 |  |  | - Turn Power to the printer on |
| 21 |  |  | - Printer will prepare ribbon and self test |
| 22 |  |  | - Close the side panel |
| 23 |  |  | - Logon to Hematrax and print full face label |
| 24 |  |  | - Keep printing labels until no visible smudges or marks are seen on the unit face label |
| 25 |  |  | - Printer is ready for use |
| Part C: Adjusting the Label Printer/Ribbon |  |  |  |
| 1 | - If the label printer does not print the label in the correct alignment |  |  |
| 2 | - Press the Setup/Exit button on the printer |  |  |
| 3 | - Hit the + button until “Label Top” is seen |  |  |
| 4 | - Hit Select |  |  |
| 5 | - Use the + or – to move the label up or down |  |  |
| 6 | - Click on Setup/Exit to save changes |  |  |
| 7 | - If label needs to be adjusted sideways… |  |  |
| 8 | - Go to Setup/Exit and select Left Position |  |  |
| 9 | - Use the + or – to move the label left or right |  |  |
| 10 | - Keep adjusting until the label is aligned and prints in the correct position |  |  |
| 11 | - Click on Setup/Exit, Save changes and exit |  |  |

References:

Hematrax ISBT 128 Blood and Blood Component Stand Alone Labeling Application, Version 6.2.1

Changing Ribbons on the ZM-400 Printer video