Purpose:

This document describes the process for reporting and tracking completion of work requests for critical equipment in order to assure appropriate revalidation after repairs before placing into service.

Procedure:

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | - Perform any troubleshooting for the equipment that is possible by staff before placing the service call. |  |
| 2 | - Complete an Equipment Out of Service Form. - Complete Notice of Laboratory Equipment Decontamination form, which is found at www.ehs.wa.edu/forms. - All equipment must be thoroughly cleaned inside and out with a disinfectant cleaner before it is sent for repair or before repair is attempted by a repair technician. | Equipment out of Service Form, Notice of Lab Equipment Decontamination Form |
| 3 | - Portable refrigerators: Notify Clinical Engineering to respond STAT (206-744-3496). - All other blood storage equipment (refrigerators, freezers, and platelet rotator): Notify Facilities Engineering to respond STAT (206-744-3191). - After hours, page the engineer on call through the hospital operator. - NOTE: Engineering is expected to respond within 30 minutes of receiving a STAT call. - Centrifuges, cell washers, plasma thawers, and small equipment are repaired by Scientific Instruments. - Main phone line 206-543-5580 - Voice Mailbox 206-685-0641 - Emergency pager 206-991-4030 |  |
| 4 | - Record the call date/time on the Equipment Out of Service Form and make a photocopy of the completed form. - The original gets taped to the equipment so it is visible to staff. - The copy goes to the supervisor or lead technologist on duty. |  |
| 5 | - When an engineer or technician arrives, record the response date/time on the Equipment Out of Service Form. |  |

| Step | Action | Related Documents |
| --- | --- | --- |
| 6 | - The Equipment Out of Service Form accompanies any equipment removed from the lab for repair. - Instruct the engineer or technician to complete the Clinical Engineering/Facilities/Scientific Instruments Sign-Off box and problem resolution field when the equipment is repaired and returned to service. | - Equipment Out of Service Form |
| 7 | - Upon completion of the repair: - Clinical Engineering OR Scientific Instruments will return the equipment with the completed Equipment Out of Service Form or - inform the lab as to when equipment repaired on site can be returned to service. - Lead Technologist or designee will review completed Equipment Out of Service Forms for completeness of documentation including action taken. - Lead Technologist and/or manager will evaluate whether revalidation is required. | - Equipment Management Policy |
| 8 | - A signature of Return to Service by Lab Staff is done to approve the work order is completed and equipment is returned or work is finished. |  |
| 9 | - If no revalidation is required, return equipment to service as follows: - For minor repairs of small equipment, do daily QC tasks and monitor under normal working situations for expected results with documentation. - For repairs of temp monitored equipment, do up to and including the quarterly QC tasks detailed in the equipment SOP usually done by lab personnel. Document in the equipment file. - Tissue storage freezers, allow to come to temp (-80˚C) and perform maintenance up to and including the semi-annual check of high and low alarms by lab staff. |  |

References:

Standards for Blood Banks and Transfusion Services, Current Edition. American Association of Blood Banks. AABB Press, Bethesda, MD.

Roback J (ed). Technical Manual, Current Edition. AABB Press, Bethesda, MD