Purpose:

This procedure provides instructions for using the Blood Product Testing (BPT) function in Sunquest, to retest blood types for red blood cell products coming into Transfusion Service inventory, converting units from unprocessed to available inventory. This SOP is to be used in conjunction with “Blood Product Entry in Sunquest”.

Method

| Step | Actions | Related Documents |
| --- | --- | --- |
| 1 | - Enter units into inventory - Complete Visual Inspection and result. - If unit passes inspection, result as ‘OK’, and proceed. - If unit does not pass inspection, quarantine the unit per SOP. | - SQ—Blood Product Entry - Visual Inspection of Red Cell Products - Visual Inspection of Plasma Products - Quarantine of Blood Products |
| 2 | - Place products in refrigerator on specified unprocessed shelves. | - Blood Product Storage Policy |
| 3 | - In small batches remove products from refrigerator. - Remove segments. - 2 segments for retention - Label with Unit number sticker - Place in dated storage bag for retention. - 1 segment for the ABORH retype. - Label with Unit number, adding a 1 or 2 if necessary for container type. - Place units back into the refrigerator. | - Sample Management Process |
| 4 | Type Confirmations on Tango |  |
| 4.1 | - Place unit sticker on test tube in vertical manner to be read by instrument. - Discharge as much blood as possible from the segment into the tube. - Place tubes into the B racks for testing on Tango. | - Tango Donor Unit Retyping |
| Step | Actions | Related Documents |
| 4.2 | - Load onto Tango. - Ensure all barcodes are read properly. - If both container 1 and container 2 products are received, then only place container 1 on Tango. Do the recheck of container 2 by hand. | - Tango Donor Unit Retyping |
| 4.3 | - Select Donor as the test to be done. This rechecks the ABO and the Rh. | - Tango Donor Unit Retyping |
| 4.4 | - Using Blood Bank Instruments function, validate results as they come off the Tango. | - Using Blood Bank Instruments Function in Tango |
| 5 | Type Confirmations—Manual Tube Method - Confirm ABO by testing all Group A and Group B units with Anti-A and Anti-B reagent. - Confirm ABO by testing all Group O units with Anti-A,B. - Confirm all RhD negative units by testing with Anti-D reagent. - Do not confirm Rh for Rh Positive units. | - ABO Typing by Tube Method |
| 5.1 | - Enter results in BPE, using the Worklist Mode. - Scan in the unit number - Click on “Add” button on screen, to create list, or use the Worklist number generated upon entry of units into inventory. - When list is complete, Click on “Continue” button on screen. - The first unit number will appear in the testing area of the screen. - Confirm unit number in gray header before entering results. - Use the Home key to get to the grid for entering ABO/RH results. - All reaction fields in the grid must be resulted. - Use ND (not done) for any tests not performed. - Enter ND for the Rh on all positive units, as Rh confirmation is not done on Rh Pos units. | - Blood Product Entry in Sunquest |
| 5.2 | - Enter single results that did not cross over from Tango to Sunquest - Using the Journal Worksheet from the Tango manually enter the results. - Enter only the testing that would be required for manual tube testing, listed in step 5 above, i.e. the TANGO does Rh on all units, but only enter the Rh results for negative units. - Add a test to the unit by tabbing to the field under ABORH and entering ;BBCS (Blood Bank Comment Suppressed) to leave a notation that the test came from Tango results. |  |
| Step | Actions | Related Documents |
| 6 | - Ensure all units have been properly typed and ready for placement into inventory. - Open the Blood Bank Inventory/Supplier Search function. - Enter H for HID and RBCG for Component Type. - In Search Mode select Unprocessed and click on “Search” - If no units are located in the unprocessed inventory, testing in completed. - If there are units remaining in unprocessed inventory, reconcile the worklist to make sure all tested units have been resulted. - Remove units from the unprocessed inventory area in the refrigerator, and place them in the available inventory area. | - Blood Product Storage Policy |

References:

AABB Standards for Blood Banks and Transfusion Services, Current Edition.

Blood Bank User Guide, Misys Laboratory