UW Medicine - Pathology

6000-02-03-02

H&E Staining Prisma Protocol

Adopted Date: 12/10/10 Review Date: Revision Date: 04/08/11

PURPOSE

To give directions for setting staining times on the Prisma H&E stainer.

SCOPE

Histology H&E staining

PROCEDURE

Tissue-Tek PRISMA H&E Staining Protocol

UWMC Histology

| STEP # | PROCESS | TIME | REAGENT BRAND |
| --- | --- | --- | --- |
| 1 | Drying | 15 min | 65 degrees (Convection Oven) |
| 2 | Xylene | 3 min | RAS* Xylene |
| 3 | Xylene | 3 min | RAS Xylene |
| 4 | Xylene | 2 min | RAS Xylene |
| 5 | Abs. Alcohol | 1 min | RAS Flex 100 |
| 6 | Abs. Alcohol | 1 min | RAS Flex 100 |
| 7 | 95% Alcohol | 30 sec. | RAS Flex 95 |
| 8 | Water Wash | 1 min |  |
| 9 | Hematoxylin | 2:45 | RAS Hematoxylin #7211 |
| 10 | Water Wash | 1 min |  |
| 11 | Acid Alcohol | 40 sec | RAS Clarifier 2 |
| 12 | Water Wash | 1 min |  |
| 13 | Bluing | 40 sec | RAS Bluing Reagent |
| 14 | Water Wash | 1 min |  |
| 15 | 95% Alcohol | 30 sec | RAS Flex 95 |
| 16 | Eosin | 45 sec | RAS Eosin Y |
| 17 | Abs. Alcohol | 1 min | RAS Flex 100 |
| 18 | Abs. Alcohol | 1 min | RAS Flex 100 |
| 19 | Abs. Alcohol | 1 min | RAS Flex 100 |
| 20 | Xylene | 1 min | RAS Xylene |
| 21 | Xylene | 1 min | RAS Xylene |
| 22 | Xylene | 1 min | RAS Xylene |
| 23 | End Station | ------- | RAS Xylene |

* RAS = Richard-Allan Scientific

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________