UW Medicine - Pathology

6000-02-07-15

Procedure for Microwave Oven Calibration

Adopted Date: April 12, 2013 Review Date: Revision Date:

PURPOSE

To describe the procedure on how to calibrate the H2800 Microwave Oven twice a year.

SCOPE

For all histotechnician and histotechnologist .

PROCEDURE

- Twice yearly, Microwave Oven needs to be calibrated to check for temperature reproducibility utilizing the base temperature chart.

- Prepare 3 containers of DI water using the plastic coplin jars with the vented covers.

- Place the container of water in the center of Microwave one at a time.

- Program time for 60 seconds and start microwave power.

- Stir and record temperature of the water.

- Repeat the testing one at a time by increasing the timing with increments of 30 seconds (60, 90, 120).

- Submerged a calibrated approved thermometer on each coplin jar right after the timer goes off. ( Be aware that some coplin jar may be hot).

- Record the temperature.

- Compare the temperature with the base temperature guide. Temperature should be within the base temperature and can only be off by a minimal degree.

- Record the temperature in the Microwave Maintenance QC log.

Acceptable Temperature Ranges

- 1 minute: 46 degrees C

- 1.30 minutes: 60 degrees C

- 2 minutes: 77 degrees C

*Notify Supervisor of any major discrepancies.

REFERENCE

RELATED DOCUMENT

APPENDIXES

Written By:					Director Approval:

_______________________	    	______________________