UW Medicine - Pathology

2000-06-01-10

pH Meter QC Procedure

Adopted Date: 06/05/09 Review Date: Revision Date: 05/03/11

PURPOSE

To monitor quality control of the Beckman 360 pH Meter in the IHC Laboratory.

PROCEDURE

Beckman 360 pH Meter:

Serial Number:  1113

UWMC Number: n/a

Instillation Date: unknown

PM Number: 21523

Preventative Maintenance Procedure:

- Daily (or before each use):

  - Calibrate pH meter once a day ,before each use.

  - Check the pH of the calibrating buffers. The pH should be accurate to ±0.02.

  - Make sure the tip of the electrode is submerged in buffer.

  - Keep all calibrating buffer solutions tightly closed.

- Annually:

  - Scientific Instruments performs preventative maintenance on the unit ever 12 months.

- For problems:

  - Refer to the owner's manual and try to troubleshoot.

  - If unable to resolve problem, contact Scientific Instruments at 543-5580.

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________

IHC Supervisor