UW Medicine - Pathology

6000-01-02-08

How to Replace a Formalin Cube in the Frozen Room

Adopted Date: 2-7-2013 Review Date: Revision Date:

PURPOSE

How to properly replace an empty formalin cube in the Mopec Formalin dispensing pump system located in frozen section room.

SCOPE

This procedure pertains to all support staff working in the frozen section room. This process should not be performed by non-support staff, including residents or attending Pathologists.

PROCEDURE

- It is mandatory for all staff to wear   proper PPE when handling Formalin.

- Push the gray button which releases the hose from the cube.

- Unscrew the male Mopec O ring from the empty formalin cube.

- Place new full cube under grossing station with the plastic round seal facing up.

- Remove the tape and white plastic seal.

- Remove the detached valve wrapped in a plastic bag.

- Unscrew the cap from steam of new formalin cube.

- Screw the cap on the old formalin cube.

- Screw the male Mopec O ring on the steam of the new formalin cube.

- Insert the female fitting attached to the hose in to the male O ring fitting until you here a snap which indicates a good seal.

- Push cube to back left corner beneath grossing station.

- Turn on formalin pump with supply hose in discard funnel and bleed out all the air in the system until you see a steady stream of formalin.

- Check the Mopec O ring for any leaks.

REFERENCE

RELATED DOCUMENT

APPENDIX

Written By:					Director Approval:

Steve Rath

______________________		       ________________________