UW Medicine - Pathology

6000-01-05-30

Band Saws Maintenance and Decontamination Procedure

Adopted Date: 01/23/12 Review Date: Created Date: 01/23/12

PURPOSE

To disinfect the bone saws on a regular basis.

SCOPE

All staff sectioning bone on the bone saws

PROCEDURE

Principle:

According to CAP regulations, the band saws should be cleaned and decontaminated with a tuberculocidal disinfectant at a time interval appropriate for the institution.  The regulations additionally state that the saw must be clearly marked as contaminated if sectioning is performed on tissue from a patient known or suspected to be positive for HIV, hepatitis B or C, SARS-related coronavirus, prion disease such as Creutzfeldt-Jakob disease, or mycobacterial or systemic fungal disease.  It must then be decontaminated before further use.

General Precautions:

- Prior to doing anything the power cord to the saw must be unplugged.

- Fluid resistant barrier gowns or lab coats, puncture and penetration resistant gloves, N-95 mask and eye protection must be worn when performing band saw disinfection procedures.  Virugard can be corrosive and can cause skin damage.

Reagents:

- ViraGuard (Veridian Corporation)

- 70% Ethanol

- 100% Ethanol

Procedure:

  - Daily Saw Cleaning for large band saws with dull blade:

    - Unplug the power cord.

    - Remove used blade if dull.  Although not a requirement, steel mesh gloves should be worn when changing the blade.

    - Dispose of blade in a provided sharps container if dull.

    - Vacuum out all debris from the inside housing and under saw with shop vac.

    - Debris must be removed because organic materials (blood and proteins) may contain high concentrations of microorganisms and could possibly inactivate the chemical disinfectant or prevent access to contaminated surfaces.  It should be treated as biohazardous waste and disposed of according to the policies and procedures of the institution.

    - Wipe down all surfaces of the saw inside and out the with  70% alcohol.

    - 70% ethyl alcohol poured on a cloth should be used to clean the saw and provide some disinfection capabilities.  Be sure to treat all surfaces that might touch fresh tissue.

  - Daily Saw Cleaning for large band saws with Sharp blade:

    - Unplug the power cord.

    - With a forceps manually rotate the blade with one hand and hold a second pair of forceps with a towel soak with 70% alcohol around the blade. Steel mesh gloves should be worn when cleaning the blade.

    - Next remove the blade and set aside.  Although not a requirement, steel mesh gloves should be worn when changing the blade.

    - Vacuum out all debris from the inside housing and under saw with shop vac.

    - Debris must be removed because organic materials (blood and proteins) may contain high concentrations of microorganisms and could possibly inactivate the chemical disinfectant or prevent access to contaminated surfaces.  It should be treated as biohazardous waste and disposed of according to the policies and procedures of the institution.

    - Wipe down all surfaces of the saw inside and out the with  70% alcohol.

    - 70% ethyl alcohol poured on a cloth should be used to clean the saw and provide some disinfection capabilities.  Be sure to treat all surfaces that might touch fresh tissue.

    - Reinstall the sharp used blade.

  - Daily Saw Cleaning for small diamond band saw :

    - Unplug the power cord.

    - Manually rotate the blade with one hand and hold  a towel soak with 70% alcohol around the blade. Although not a requirement, Steel mesh gloves could be worn when cleaning the blade.

    - Next remove the blade and set aside.  Although not a requirement, steel mesh gloves could be worn when removing the blade.

    - Vacuum out all debris from the inside housing and under saw with shop vac.

    - Debris must be removed because organic materials (blood and proteins) may contain high concentrations of microorganisms and could possibly inactivate the chemical disinfectant or prevent access to contaminated surfaces.  It should be treated as biohazardous waste and disposed of according to the policies and procedures of the institution.

    - Wipe down all surfaces of the saw inside and out the with 70% alcohol.

    - 70% ethyl alcohol poured on a cloth should be used to clean the saw and provide some disinfection capabilities.  Be sure to treat all surfaces that might touch fresh tissue.

    - Reinstall the sharp used blade.

  - Chemical Disinfection / Decontamination of both saw types:

      - Disinfection / Decontamination should be performed at an interval deemed appropriate with the usage of the saw.

        - If the saw is used daily, it must be decontaminated weekly.

          - Pour disinfectant (Viraguard) onto surfaces or absorbent disposable towels and wipe the contaminated surfaces and leave wet for 10 minutes.

          - All accessories must be taken apart, disinfected with Viraguard and dried thoroughly.

          - Do not create aerosols by spraying disinfectant (or anything else) in an open saw chamber.

          - Any disposable material used in the disinfection process must be disposed of in accordance with the Compromised Specimen procedure, 6000-01-04-01, and Safety, Infection Control Management policy, 100-05-01-01.

  - Following Disinfection:

            - After the disinfection procedure is complete, use 70% alcohol and wipe down all surfaces.

            - Five cc of liquid bleach should be add to the liquid waste container in the smaller wet diamond band saw.

  - Maintenance:

            - Examine the blades and moving parts for defects.  Notify the maintenance department if repairs are needed.

Note:  The germicidal activity of ethyl alcohol is most effective in the 70% range because it can penetrate tubercle bacteria and it has an advantage over isopropyl alcohol of being able to kill hydrophilic viruses.

Note: The saws must be decontaminated with a tuberculocidal disinfectant at an interval appropriate for the institution; this must be a weekly for saw that are used daily.  Trimming and sections of tissue that accumulate inside the saws must be removed during daily cleaning.  Although not a requirement, steel mesh gloves should be worn when changing knife blades.

REFERENCE

- Carson, Freda, Histotechnology a Self-Instructional Text 2nd edition, ASCP Press.  Chicago, IL ASCP Press, 1990.

- ViraGuard MSDS

Written by:					Input/Reviewed by:

Steven Rath M.H.S. P.A. (ASCP)		Dan Luff, PA

UWMC Pathology Medical Director Approval:

_______________________________________