Purpose:

This procedure provides instructions for performing reagent QC for manual testing.

Policy:

- All antisera and cellular reagents are stored in a monitored refrigerator at 2-8°C.

- In use reagents are arranged in numbered racks for manual testing.

- Racks are rotated in alphabetical order ( A, B, C, A, etc.)  for Daily Quality Control testing.

- One rack is used each day for Daily Quality Control testing.

- Each lot of reagents is tested before being placed into use.

| Reagents Requiring Daily QC Testing |  |
| --- | --- |
| Routine Reagent Rack - anti-A - anti-B - anti-A,B - A1 cells - B cells - anti-D - Screening Cells 1, 2, 3 - LISS - anti-IgG - antihuman globulin (AHG) control cells - Seraclone ABO + Rh Control | Direct Antiglobulin Reagent Rack - Polyspecific antihuman globulin |

Sources of Control Materials for Antibody Detection and ABO Reverse Testing:

| Antibody QC | Source | Antibody |
| --- | --- | --- |
| Primary | Ortho Confidence QC Kit | Anti-A, -B, -D, -c |
| Secondary | BioRad Solid Screen II Control B | Anti-c |
|  | BioRad Solid Screen II Positive Control | Anti-D |
|  | TANGO QC Controls | Anti-A, -B |
| Option requiring TS Manager Approval | Patient source of antibody(ies) | Must react with all antibody detection cells |

Sources of Control Materials for ABO and D Antigen Forward Testing:

| Antigen QC | Source | Antibody |
| --- | --- | --- |
| Primary | Ortho Confidence QC Kit | Cell 1: A1B rr Cell 2: O R1r |
| Secondary | TANGO QC Controls | Combination of types to test anti-A, -B, and –D reactivity and specificity |
| Option requiring TS Manager Approval | Other source of antigen(s) | Combination of types to test ABO and D antisera reactivity and specificity |

Procedure:

|  | Action | Related Documents |
| --- | --- | --- |
| 1 | - Examine reagents for color and appearance: - In date - Clear and no turbidity in antisera - No Hemolysis in reagent red cells - Reagent color has not changed - Reagents and QC must meet criteria before testing is performed. |  |
| 2 | - Record testing date, lot number and expiration date of reagents. | - Daily Reagent QC Form |
| 3 | - Revise Daily QC Form as required for additional reagent QC: Example: PEPG QC’d in addition to LISS QC | - Daily Reagent QC Form |
| 4 | - Label tubes as per in table 1 | - Table 1 |
| 5 | - Perform testing per applicable procedures for the test being controlled. |  |
| 6 | - Record Rack number and test result on Daily QC Form |  |
| 7 | - Interpret test results and evaluate for acceptability. | - Reading and Grading Tube Reactions |
| 8 | - Record Tech ID. |  |
| 9 | - Obtain 2nd Tech review |  |
| 10 | - Perform problem resolution for any unacceptable results as follows: - Repeat testing on original bottle of reagent. - If retest is unacceptable, test on a new lot number of same reagent. - Quarantine unacceptable lots and notify TSL manager. - Complete QIM to document. | - Reagent Receipt and Discard |

Table A: Tube Labeling:

| Tube | Contains | Contents | Expected Results |  |  | Related Document |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Anti-A | Ortho Confidence Cell 1 (A1B Neg rr) or Acceptable QC source | 2-4+ |  |  | - ABO Rh Tube Method |
| 2 | Anti-B |  | 2-4+ |  |  |  |
| 3 | Anti-A,B |  | 2-4+ |  |  |  |
| 4 | Anti-A | Ortho Confidence Cell 2 (O Pos-R1r) or Acceptable QC source | 0 |  |  |  |
| 5 | Anti-B |  | 0 |  |  |  |
| 6 | Anti-A,B |  | 0 |  |  |  |
| 7 | Anti-D |  | 2-4 |  |  |  |
| 8 | Anti-D | Ortho Confidence Cell 1 (A1B Neg rr) or Acceptable QC source | IS/IgG=0 CC=2-4+ |  |  |  |
| 9 | A1Cells | Ortho Confidence Antibody Reagent (Anti-A, Anti-B, diluted 1:3) or Acceptable QC source | 2-4+ |  |  |  |
| 10 | A2 Cells |  | 2-4+ |  |  |  |
| 11 | B Cells |  | 2-4+ |  |  |  |
| 12 | A1Cells | Seraclone | 0 |  |  |  |
| 13 | A2 Cells |  | 0 |  |  |  |
| 14 | B Cells |  | 0 |  |  |  |
| 15 | Coombs Control cells | AHG Poly | 2-4+ |  |  | - DAT by Tube Method |
| 16 |  | Anti-C3 | 0 |  |  |  |
| 17 | Complement coated Control cells | AHG Poly | 1-4+ |  |  |  |
| 18 |  | AHG IgG | 0 |  |  |  |
| 19 |  | Anti-C3 | 1-4 |  |  |  |
| POSITIVE ABS LISS |  | Positive Control: Ortho Confidence Antibody Reagent (Anti-A, Anti-B, Anti-c, Anti-D, diluted), or Acceptable QC source Negative Control: Seraclone | 37 | IgG | CC | - Antibody Screen by LISS Tube Method |
| 20 | SC I |  | 0 | 1-4 | NT-4+ |  |
| 21 | SC II |  |  |  |  |  |
| 22 | SC III |  |  |  |  |  |
| Negative ABS LISS |  |  | 0 | 0 | 2-4+ |  |
| 23 | SC I |  |  |  |  |  |
| 24 | SC II |  |  |  |  |  |
| 25 | SC III |  |  |  |  |  |

References:

Standards for Blood Banks and Transfusion Services, Current Edition, Bethesda, MD: American Association of Blood Banks.

Roback J (ed). Technical Manual, 16th Edition. AABB Press, Bethesda, MD. 2008.

Current version of reagent manufacturer’s package insert instructions

CAP Accreditation Program, Transfusion Medicine Checklist, Current Version