Purpose

To describe the quality control and maintenance procedures performed, ensure blood components that require cold storage are maintained within regulatory temperature guidelines and required documentation is available for review.

| Step | Action | Related Documents |
| --- | --- | --- |
| Monthly Maintenance |  |  |
| Temp Comparisons and Alarm Response Check | - Compare Temp Trak readings against manual temperature readings in each unit: - Record manual temperature reading top and bottom - Record digital reading - Record simultaneous Temp Trak reading - Temperatures must agree within 2o degrees C. - If greater than 2oC, go to section “Troubleshooting”. - Perform Audible and Temp Trak alarm testing: - Select either HIGH or LOW alarm activation - Change Temp Trak probe reading interval to 5 minutes. - Utilizing a warm or cold solution, placed as close to or around the probe as possible, cause the audible alarm and pager Temp Trak alarm to activate. - If Temp Trak pager did not respond when audible alarm sounded, maintain probe in solution until the Temp Trak pager sounds. - Record on Monthly Storage Unit Alarm Log: - Unit number, date and time - Audible alarm sounded? Y / N - Top, bottom and digital temperatures - Temp Trak temperature and Y / N when pager is activated - Tech ID - Change Temp Trak probe reading interval to 15 minutes. | Using the Temp Trak System to Monitor Blood Storage Devices Monthly Storage Unit Temperature Log Monthly Alarm Activation Record |

| Step |  | Action |  | Related Documents |
| --- | --- | --- | --- | --- |
| Quarterly Maintenance |  |  |  |  |
| HVAC | Schedule HMC HVAC - Check door hinges - Check door gaskets - Vacuum condensing units |  | Daily Equipment, Temperature, and Inventory Duties Checklist |  |
| High/ Low Alarm Testing | High and low alarms are tested. Newer refrigerator and freezer models can be done electronically; older models will be tested with a manual method described in detail in steps to follow. Temp Trak alarms will be verified quarterly. - Place the alarm probe and a NIST calibrated thermometer in a full container of glycerol or similar solution and verify the starting temperature is between 1-6°C. - Low alarm activation - Perform first. Place the container with the probe and thermometer in a pan containing ice/water slush -1°C or colder. - Add several spoonfuls of table salt to help lower the temp. - Close the refrigerator door. - Agitate the container periodically until the alarm sounds. - Record the temperatures as the low alarm (chart, digital, internal, Temp Trak). - Silence the audible alarm. - High alarm activation - Repeat sequence for substituting a pan of cool water for the ice/water slush that is a temperature of 12-15°C. No salt needed. - Record the temperatures as the high alarm (chart, digital, internal, Temp Trak). - Silence the audible alarm. - Respond to Temp Trak pager alarm: High and Low - If Temp Trak pager did not respond when audible alarm sounded, maintain probe and NIST in solution until the Temp Trak pager sounds. - See Troubleshooting section if temperatures are not consistent with expected alarm settings |  | Refrigerator/ Freezer Maintenance Form Using the Temp Trak to System to Monitor Blood Storage Devices |  |

| Step | Action |  | Related Documents |
| --- | --- | --- | --- |
| Troubleshooting |  |  |  |
| Temperature readings are not within the acceptable limits or do not agree within + 2˚C | Circle the unacceptable temperature reading(s) and begin corrective action. - Verify plugs are securely seated in the outlets. - Open and close door firmly, wait 10 minutes and recheck the temperature. - Record follow up temperature separated by forward slash and document the corrective action taken in the comments field. - If temperatures do not come back into acceptable limits within 15 minutes, then remove equipment from service. | Removing Equipment From Service Equipment Out Service Form Refrigerator and Freezer Maintenance form |  |
| Alarm does not sound | - Slide the digital monitor box up and release it from the anchor. - Check the battery to ensure that the contact wires are snapped firmly into place. - Try alarm again. - If still it doesn’t sound, then replace the 9-volt battery. - Try alarm again. - If the alarm still will not sound, remove equipment from service. |  |  |
| LED Status is flashing green | - Check power source and replace the 9-volt battery if necessary. |  |  |
| Temp Trak | If Temp Trak is not responding or temperature differences indicate the Temp Trak is incorrect: - Contact TempTrak at 1-888-533-6900 |  |  |

References:

Standards for Blood Banks and Transfusion Services, Current Edition. American Association of Blood Banks. AABB Press, Bethesda, MD

Helmer i Series Refrigerator Service Manual

Helmer i Series Freezer Service Manual

Jewett Refrigerator Service Manual