Purpose

To describe the quality control and maintenance procedures performed to ensure autologous bone specimens are maintained within specified temperature guidelines and required documentation is available for review.

| Step | Action | Related Documents |
| --- | --- | --- |
| Monthly Maintenance |  |  |
| Temp Comparisons and Alarm Response Check | - Compare Temp Trak readings against manual temperature readings in each unit: - Record manual temperature reading top and bottom - Record digital reading - Record simultaneous Temp Trak reading - Temperatures must agree within 2o degrees C. - If greater than 2oC, go to section “Troubleshooting”. - Perform Audible and Temp Trak alarm testing: - Perform HIGH alarm activation - Change Temp Trak probe reading interval to 5 minutes. - Utilizing a warm solution, placed as close to or around the probe as possible, cause the audible alarm and pager Temp Trak alarm to activate. - If Temp Trak pager did not respond when audible alarm sounded, maintain probe in solution until the Temp Trak pager sounds. - Record on Monthly Storage Unit Alarm Log: - Unit number, date and time - Audible alarm sounded? Y / N - Top, bottom and digital temperatures - Temp Trak temperature and Y / N when pager is activated - Tech ID - Change Temp Trak probe reading interval to 15 minutes. | Using the Temp Trak System to Monitor Blood Storage Devices Monthly Storage Unit Temperature Log Monthly Alarm Activation Record |

| Step | Action | Related Documents |
| --- | --- | --- |
| Monthly Maintenance |  |  |
| Clean | - Perform monthly at a minimum; more often as needed - Clean door gaskets, hinges, inner doors - Using a soft cloth, remove any frost build-up form the gasket and doors. - Clean probe cover. - Remove any ice build-up around sensor probe. - Clean Vacuum Relief Port - Observe the inner side of the port periodically for frost and ice build-up and remove any frost with a soft dry cloth. - Clean Air Filter. - Open the front lower door by grasping the bottom left corner. - Locate the grille on the door. Grasp the middle of the grille material and gently pull out to remove. - Wash the filter material using water and a mild detergent. - Dry by pressing between two towels - Install the filter back into the grille and attach the grille. - Check System Battery - Press the Mode key until Configuration indicator lights. - Press and right arrow until SYS BAT TEST is displayed in the message center - Press Enter to initiate the test. - TESTING BATT will display during the testing period. - At completion BATT GOOD or BATT FAIL will display. If test fails replace battery. | Refrigerator and Freezer Maintenance Form |

| Quarterly Maintenance |  |  |  |
| --- | --- | --- | --- |
| Alarm Testing | High Alarm Test - Press the Mode key until Configuration indicator lights. - Press the right arrow until HI ALRM TEST is displayed in the message center. - Press Enter to initiate the test. - The temperature on the display will begin to increase until the high alarm set point (-65°C) has been reached. - The audible alarm will sound and the alarm indicator will flash. - Press the Mute key to silence the alarm. Low Alarm test - Press the Mode key until Configuration indicator lights. - Press the right arrow until LO ALRM TEST is displayed in the message center. - Press Enter to initiate the test. - The temperature on the display will begin to decrease until the high alarm set point (-90°C) has been reached. - The audible alarm will sound and the alarm indicator will flash. - Press the Mute key to silence the alarm. - If Alarm does not sound take equipment out of service |  | Refrigerator and Freezer Maintenance form |
| Troubleshooting |  |  |  |
| Temperature readings are not within the acceptable limits or do not agree within + 2˚C |  | - Circle the unacceptable temperature reading(s) and begin corrective action. - Open and close door firmly, wait 10 minutes and recheck the temperature. - Record follow up temperature separated by forward slash and document the corrective action taken in the comments field. - If temperatures do not come back into acceptable limits within 15 minutes, then remove equipment from service. - Relocate product to acceptable storage unit. - Notify Engineering - Document all decisions and new location of bone specimens. | Removing Equipment From Service Equipment Out Service Form Refrigerator and Freezer Maintenance form |
| Temp Trak |  | If Temp Trak is not responding or temperature differences indicate the Temp Trak is incorrect: - Contact TempTrak at 1-888-533-6900 |  |

References:

Standards for Blood Banks and Transfusion Services, Current Edition. American Association of Blood Banks. AABB Press, Bethesda, MD

VWR Model 5600 Series -86 ULT Freezer Manual