Purpose:

This procedure describes the daily maintenance to be performed on the TANGO instrument to assure that it is maintained properly according to the manufacturer’s recommendations.

Required Materials:

- Tango Maintenance Form

- Additional TANGO Reagent Form

- Lint free wipes

- DI water

- Disinfectant

Procedure:

| Step | Action | Related Documents |
| --- | --- | --- |
| General Statements: |  |  |
| 1 | Use only the Daily Maintenance section of the Tango Maintenance Form and complete the Additional Tango Reagent Form in its entirety | Tango Maintenance Form Additional TANGO Reagent Form |
| 2 | Record a check mark as each task is completed and enter tech intials in the last box. |  |
| Rinsing Pipettor Needles/Cleaning Washer Needles |  |  |
| 3 | The TANGO can be automatically programmed to perform this function. - Select the ‘Auto Rinsing’ button to open the ‘Automatic Rinsing’ window. - Define the intervals for automatic rinsing. - For automatic rinsing, the system automatically switches over to the cleaning solution in the Cleaning Solution container. - It is recommended that the Auto-Rinsing repetition be set to once a day at a time that is convenient to workflow. - Select the ‘Active’ button so that the check mark appears in the box. This activates the Auto Rinsing function. |  |
| 4 | View Event Log in the lower left hand corner of the screen to verify auto rinsing took place. Alternately, the presence of a checkmark indicating the rinse is “Active” may be verified in the Auto Rinse dialog. |  |
| Step | Action | Related Documents |
| Checking the Suspension Cup |  |  |
| 5 | If the instrument is turned off, remove the acrylic glass cover. If the instrument is in ‘Ready’ status, follow Removing the Covers SOP to remove the cover. | Removing the Covers |
| 6 | Remove the gray plastic cover. |  |
| 7 | Without removing the cup, verify it is empty and inspect for cracks, liquid or residue. |  |
| 8 | If the cup is dirty or cracked, it needs to be replaced. Follow Monthly Maintenance SOP. | Monthly Maintenance |
| 9 | The plastic splash bowl should be dry or contain only a small amount of residual liquid. |  |
| 10 | Replace gray plastic cover. Make sure the gray plastic cover is in line with the TANGO housing |  |
| 11 | Reinsert acrylic glass cover. |  |
| Checking tubing line for bubbles |  |  |
| 12 | Observe tubing that runs from the top of the analyzer to the pipette needles. - Small bubbles are normally present in the tubing. - Large bubbles indicate leaking syringes or empty containers. |  |
| 13 | Check fluid containers to ensure there is sufficient fluid. Empty waste container as necessary. | Managing Liquid Containers |
| 14 | Check syringes for leaking. If leaking is observed or bubbles persist, call Technical Service. |  |
| Cleaning the TANGO surfaces |  |  |
| 15 | Clean up any splatters/spills immediately. |  |
| 16 | The surface of the TANGO can be disinfected with a suitable disinfectant. |  |
| 17 | After disinfection, the surfaces should be wiped again with DI water to remove any residual disinfecting solution. |  |
| Checking Pipettor needles |  |  |
| 18 | Check the pipettor needles for damage. The pipettor needles should be observed during pipetting for the following: - The needles should be straight. - The flow of water during washing should be straight down. - There should be no dripping from needles after initialization or auto-rinsing. - Liquid should flow through the needles during washing. |  |
| 19 | Gently wipe the exterior of the needles with a lint free wipe moistened with DI water. |  |
| Document Reagent Lot Numbers in Use |  |  |
| 20 | The on-board reagents can be printed by selecting a CID of a specimen with current date. - In main menu select Lab Journal and enter CID of specimen in search box. - Touch specimen selected to highlight and then Export. - In the Export screen select the report for Lot Information by Date. Date should be current date. - Select Export key to get printout. |  |
| 21 | Other lot numbers in use will need to be filled out on Additional Tango Reagent Form | Additional TANGO Reagent Form |
| 22 | Perform Reagent Dating Verification, in the appropriate column: - Record open/draw date and expiration date/time from the reagent bottle or sample label. - Mark YES or NO for acceptability |  |
| 23 | Add to the form throughout the day as lot numbers are opened or patient samples are selected as controls. |  |
| 24 | Review Reagents for Acceptability and Performance. - Document with Two Tech ID sign-off and Date. - File in Tango Daily QC binder. - Tango user on every shift will perform a daily review - TS Lead or designee will perform a daily review. - TS Manager performs a monthly review. |  |

References:

TANGO User Guide, Version 3.3