# POC.35 ACT Kaolin # Using the i-STAT® 1System

| Prepared By | Date Adopted/Replaces | Approved By |
| --- | --- | --- |
| Stacy Reed/Trina Kimbrough | 06/29/2017 | Edward H. Adelstein, M.D. |

| REVIEW |  |  |  |
| --- | --- | --- | --- |
| Review or Revision Date | Comments (check all that apply) |  | Approved By |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
|  | No changes Revision initiated – see below Available on-line |  |  |
| Date | Page # | Description of revision |  |

| Distribution |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Hospital Intranet |  |  |  |  |  |  |
| ATC Office |  |  |  |  |  |  |

| Retirement Date | Approved By | Archived Location |
| --- | --- | --- |

# POC.35  ACT Kaolin

Using the i-STAT® 1System

INTENDED USE

The i-STAT Kaolin Activated Clotting Time (KaolinACT) test is an in vitro diagnostic test that uses fresh, whole blood, and is used to monitor high-dose heparin anticoagulation frequently associated with cardiovascular surgery.

PRINCIPLE

The i-STAT® Kaolin Activated Clotting Time test, KaolinACT, is a measure of the time required for complete activation of the coagulation cascade.

In traditional ACT tests, coagulation is initiated by mixing a whole blood sample with a particulate activator, and complete activation is indicated when extensive or localized clots form as activated thrombin converts fibrinogen to fibrin. These clots are mechanically detected.

The i-STAT KaolinACT test is similar to traditional ACT tests except that the endpoint is indicated by the conversion of a thrombin substrate other than fibrinogen and an electrochemical sensor is used to indicate the event of this conversion. The substrate used in the electrogenic assay has an amide linkage that mimics the thrombin-cleaved amide linkage in fibrinogen.

The substrate is H-D-phenylalanyl-pipecolyl-arginine-p-amino-p-methoxydiphenylamine which has the structure:

Phenylalanine - Pipecolic acid - Arginine -- NH - C6H4 - NH - C6H4 - OCH3

Thrombin cleaves the amide bond at the carboxy- terminus of the arginine residue (denoted by the two dashes) because the bond structurally resembles the thrombin-cleaved amide linkage in fibrinogen. The product of the thrombin-substrate reaction is the electrochemically inert tripeptide Phenylalanyl - Pipecolyl - Arginine and the electroactive compound NH3+ - C6H4 - NH - C6H4 - OCH3. The formation of the electroactive compound is detected amperometrically, and the time of detection is measured in seconds. The test reports the Activated Clotting Time (ACT) in seconds.

CLINICAL SIGNIFICANCE

The ACT cartridge as part of the i-STAT system is a quantitative assay for monitoring moderate and high-level heparin therapy through analysis of arterial and venous whole blood samples. The ACT is primarily used to monitor a patient’s state of anticoagulation due to heparin administered during a medical or surgical procedure (e.g. cardiac catheterization, renal dialysis, Percutaneous Transluminal Coronary Angioplasty).

SPECIMEN COLLECTION AND HANDLING

Use standard biohazard precautions when collecting any blood specimen.  Wear gloves during specimen collection and while testing.  Dispose of needles in appropriate puncture resistant containers.

The I-Stat ACT test can be performed using venous or arterial samples. No Anticoagulants.

Specimen Labeling

The patient is identified by the patient ID bracelet, by asking patient to state their name and SSN, or by hospital “time out” policy performed in OR settings in lieu of asking the patient to respond. Because the specimen is analyzed immediately after collection; tested in the same room where the patient procedure is being performed; tested by one authorized user, and then discarded, there are no requirements for specimen labeling.  In the event the specimen is taken into another procedure room or testing performed by a second authorized user, the specimen container must be labeled with the patient’s full name and full SSN.

Collection

If a syringe is used, it should have a 23 gauge needle or larger.  DO NO use excessive force when expelling the blood specimen through the needle.  This may cause hemolysis.

Blood Samples to be used for coagulation testing must be collected according to the following procedures to assure the integrity of the fresh whole blood sample.

Syringe sample, from indwelling line:

NOTE: The amount of blood required to adequately flush the line until it is free of contaminants is dependent on the amount of solution contained within the line. A typical heparin lock will require approximately 5 cc to clear the line. Greater volumes will be required to clear longer lines.

1.	After flushing the sampling port, obtain at least 40 µl blood sample in a non-anticoagulated plastic syringe.

2.	Immediately dispense one drop of fresh whole blood into the sample well of the test cuvette, filling the well from the bottom up.  This may be done either with or without a transfer needle.  A sufficient quantity of blood must be added directly to the center of the sample well to fill it flush to the top.

## Syringe sample, from a venipuncture

- Prepare the venipuncture site by cleansing with alcohol and allowing to air dry completely.

- Obtain a minimum of 40 µl sample with the second syringe.

- Immediately dispense fresh whole blood into the sample well of the test cuvette, filling to the fill line.  This may be done either with or without a transfer needle.  A sufficient quantity of blood must be added directly to the sample well being careful not to over- or underfill the port.

- Storage conditions, transportation, and processing are not applicable since testing must commence immediately.

Suitable Specimens for ACT:

- Sample requirement is 40 µl of fresh whole blood (venous or arterial) without anticoagulant without anticoagulant, clot activators, or serum separators collected in a plastic syringe properly

- The collection device cannot contain anticoagulants such as heparin, EDTA, oxalate, or citrate.

- The collection device cannot contain clot activators or serum separators.

- The sample should be immediately dispensed into the sample well of a cartridge.

- If from an in-dwelling line, possible heparin contamination and specimen dilution should be considered.  Flush the line with 5 mL saline and discard the first 5 mL of blood or three to six dead space volumes of the catheter.

- Withdraw the sample for testing into a fresh plastic syringe.

- Device used to transfer sample to cartridge must be plastic.

- If a repeat measurement is needed, a fresh sample must be obtained

Criteria For Specimen Rejection

- Evidence of clotting

- Specimens collected in vacuum tubes with anticoagulant

- Specimens collected in glass syringes or tubes or with anticoagulant of any kind

- Other sample types such as urine, CSF, and pleural fluid

Precautions: Avoid the Following Circumstances

- Drawing a specimen from an arm with an I.V.

- Stasis (tourniquet left on longer than one minute before venipuncture)

- Extra muscle activity (fist pumping)

- Hemolysis (alcohol left over puncture site, or a traumatic draw)

- Icing before filling cartridge

- Time delays before filling cartridge

REAGENTS - SUPPLIES - EQUIPMENT

i-STAT® 1 Analyzer

i-STAT® ACT-Kaolin Cartridge

Store cartridges at 2 to 8°C (35-46°F)

Bulk storage in Pathology, room D120 refrigerator

Refrigerated cartridges are stable until manufacturer’s expiration date. Kaolin Cartridge must be at room temperature prior to use. For cartridges stored at refrigerated temperature, this may take up to 60 minutes per box or 5 minutes per cartridge.

Room Temperature Storage

Cartridges are stable for 14 days at room temperature.

Record room temperature expiration date on the outer box.

Keep all cartridges in the box until time of use to protect the inner cuvette

pack.

Contents

Each i-STAT Kaolin ACT cartridge provides a sample collection chamber, sensors to detect the coagulation endpoint, and dry reagents necessary to initiate and allow coagulation. Stabilizers and reagents are coated on a section of the sensor channel and include the following reactive ingredients:

| Reactive Ingredient | Minimum Quantity |
| --- | --- |
| Kaolin | 23.4 μg |
| Thrombin Substrate | 0.09 μg |

INSTRUMENT OPERATIONS

Start Up

Press to turn on Handheld.

PATIENT TESTING

Specimens will be analyzed immediately after collection.  Only fresh whole blood collected in plastic syringes or tubes without anticoagulant may be used for the ACT test. If samples are collected in plastic tubes, use a plastic syringe to transfer sample to the testing device. Testing must be performed on the i-STAT System with no delays

Cartridge Handling

- Do not remove the cartridge from its protective pouch until it is at room temperature (18-30°C or 64-86° F).

- Use a cartridge immediately after removing it from its protective pouch.  Prolonged exposure may cause a cartridge to fail a Quality Check.

- Do not contaminate the contact pads with fingerprints.

- Do not apply pressure to the central area of the label as the calibrant pack inside could burst prematurely.

- Do not block the air vent as the sample will not flow to the fill mark and the calibrant solution will not flow to the sensors.

- Do not use a cartridge on which blood or any other fluid has spilled.  Avoid filling cartridges on surfaces that may cause the cartridge to pick up fibers, fluid, or debris that may lodge in the analyzer.

Filling and Sealing Cartridge

- Place cartridge on a flat surface.

- Direct the tip of the dispensing device into the sample well.

- Dispense sample slowly and steadily until it reaches the fill mark indicated on the cartridge label.  Leave some sample in the sample well.

- Press the rounded end of the closure until it snaps into place.  Slightly lift finger or thumb and ensure that the cartridge is closed before completely removing finger or thumb from the closure.

Inserting and Removing the Cartridge from the Analyzer

- Inserting

  - Align the cartridge with the contact pads facing up and toward the cartridge port

  - Push the cartridge slowly and smoothly into the cartridge port until it clicks into place.

- Removing

  - Do not attempt to remove the cartridge while the message “Cartridge Locked” remains on the screen.

  - When the results are displayed, pull the cartridge straight out of the analyzer.

- Dispose of the cartridge in a container for biohazards.

Performing Patient Analysis

- Press the On/Off button

- Press 2, for i-STAT Cartridge

- Press SCAN and hold until the operator barcode is read, or manually enter your operator ID

- Press SCAN to enter the patient ID from the wrist band, or manually enter the patient ID

- Repeat patient ID entry to verify correct entry, if necessary.

- Press SCAN and hold to enter the cartridge lot number

- Unwrap the cartridge, touching only the sides of the cartridge

- Fill the cartridge according to instructions above in “Filling and Sealing Cartridge”.

- Push the sealed cartridge into the handheld port until it clicks into place.

- Wait for the test to complete.

- Review results

- Follow the screen prompt and remove the cartridge.  Discard the cartridge in a bio-hazard bag.

- Place i-Stat into downloader/recharger to transmit data to VISTA/CPRS

Shutdown

Press to turn off Handheld.

REPORTING OF RESULTS

Results are reported in CPRS (with appropriate reference ranges if applicable)

Normal range is not reported with the ACT result since each patient’s individual therapeutic range will be determined by his/her physician according to the treatment protocol for the medical procedure being performed (ie; cardiac cath, OR)

The nurse performing the test and Physician performing the procedure are responsible for proper evaluation and documentation of results and treatment necessary for the patient.

Reference Ranges are not always appropriate for Activated Clotting Time, instead they have Target Ranges that vary with procedures and protocols. In some cases, physician clinical preferences may be appropriate. As a guideline the following may be used:

Reference Ranges

Non-heparin target range < 200 sec

Heparinized target range >250 sec

INTERPRETATION OF RESULTS

Interpretation of results is at the discretion of the physician performing the procedure and based upon accepted published data.  Absolute results are not used to assess the clinical condition of the patient; therefore normal ranges will not be reported with results.  The authorized operator is responsible for proper evaluation, documentation of results, and treatment necessary for the patient.

Critical Results

Critical results are test results that fall outside high and low critical limits that define the boundaries of life-threatening values for a test. Critical results represent an emergency condition and must be reported immediately to the patient’s attending physician or nurse. Because patient testing occurs at the bedside and physician decisions are made and actions are executed based upon therapeutic ranges, there are no critical result for ACT+.

TRANSMITTING INFORMATION

- All meters should be returned to the base unit upon completion of quality control and patient testing.

- Place the meter back in the base unit and verify that the screen displays “Communication in Progress” message.  This verifies patient results are transmitting to the computer network server and information from the RALSWeb3 database is being returned to the meter.

QUALITY CONTROL

Liquid QC

ACT Kaolin

I-STAT ACT controls    Level 1 and Level 2

Procedure for testing cartridges with I-STAT Level 1 and Level 2 ACT Controls

- Prior to use, allow one vial each of the lyophilized plasma and calcium chloride reconstituting fluid to stand at room temperature for a minimum of 45 minutes.

- Remove the cap and stopper from the vials and pour the entire contents of the calcium chloride vial into the lyophilized plasma vial.  Place the stopper back on the reconstituted vial.

- Allow the vial to sit for 1 minute and then mix the contents by swirling gently for 1 minute, then inverting slowly for 30 seconds.

- Use a plastic pipette, syringe, or capillary tube without anticoagulant to transfer the solution to an ACT cartridge.

- Immediately seal the cartridge and insert it into a handheld.  This process must be completed within 30 seconds of the complete reconstitution of the control sample.

- Compare results to the value assignment sheet ranges.  If results are within the expected ranges, use the cartridges as needed.

Store at 2 to 8°C (35-46°F)

Target values are printed on a Value Assignment Sheet posted on the APOC website at www.abbottpointofcare.com.  Ensure that the lot number printed on the Value Assignment Sheet matches the lot number on the label of the ampule and that the software version above the target value table matches the software version in the handheld.

The ranges displayed represent the maximum deviation expected when controls and cartridges are performing properly.  Should results outside these ranges be obtained, refer to the Troubleshooting section that follows the Procedure for Testing Controls in the System Manual. If the repeated QC is still outside the acceptable range, the analyzer may not be used until the problem is resolved. A corrective action form will be filled out documenting actions taken for the resolution of the problem.  If the problem is not resolved, contact iSTAT technical service.

Performing Internal Electronic Simulator Test

The internal Electronic Simulator test cycle is automatically activated when a cartridge is inserted after the customized interval is reached.  If the analyzer passes the simulator test, the cartridge test cycle proceeds.  If not, the analyzer displays “ELECTRONIC SIMULATOR FAIL”.    The same cartridge can be re-inserted immediately after the FAIL message is displayed.  If the analyzer fails the simulator test again, contact the Ancillary Testing Office at extension 53026 or 53034 for assistance.

Performing External Electronic Simulator Test

This process will be completed by the Ancillary Testing Coordinator or designee.  The External Electronic Simulator and i-STAT® 1 System Manual will be stored in the Ancillary Testing Office.

Checking Thermal Probes

This process will be completed by the Ancillary Testing Coordinator or designee and by using the instructions in the i-STAT® 1 System Manual.  Thermal Probes will be checked every six months.

Performing Internal QC

The i-STAT has a lock-out in place to automatically run the internal electronic QC every 8 hours of patient testing. The analyzer will not proceed with patient testing if the internal QC does not pass. Internal QC is stored in RALS.

Any QC failure and corrective action taken is recorded on the corrective action form.

###### Proficiency Testing

Proficiency testing will be used as a system to compare the accuracy of the testing procedure and the proficiency of the users with other laboratories.  All iSTAT instruments will be tested with the CAP survey material and results reported.

All authorized operators regularly using the iSTAT instrument will complete testing of the CAP survey.  Survey samples will be rotated among staff as evenly as possible with consideration to patient workload at the time.

CALIBRATION

Calibration of the i-STAT® system is automatically performed as part of the test cycle on each cartridge and does not require any action by the operator.

The ACT cartridge is self-calibrating. Each cartridge contains a known concentration of the analyte. The signals produced by the sensors in response to the calibrant are compared to the sensors response to the sample. No other calibration is needed per manufacturer’s instructions.

Linearity

College of American Pathologist accreditation requires the use of a 3-point calibration including a minimum, mid-point, and maximum values for calibration verification. This requirement is non-applicable for coagulation assays. Therefore, linearity materials will be used to verify calibration after major maintenance or service, or at least once every six-months.

CALIBRATION VERIFICATION

Calibration verification is a procedure intended to validate instrument performance according to manufacturer expectations. Using the i-STAT System for nonwaived testing and following manufacturer’s recommendations, calibration verification will take place for the following:

- Before any new i-STAT System is placed into use.

- At changes of reagent lots upon initial use of the i-STAT System until lot-to-lot variation is proven to be consistent.  Verify the new lot using the Calibration Verification Kit for the first six lot numbers comparing each to the previous lot(s).  Lot to lot variation must be within the evaluation limit of +0.3 mg/dL or 15%, whichever is greater.

- To verify that results have not been affected by major maintenance or repair.

- To troubleshoot when control values are out-of-range or reflect an unusual trend or shift.

- As recommended by manufacturer

- At least every six months

Calibration is performed every six months by testing all five levels of i-STAT TriControls Calibration Verification material in duplicate.   Target values are obtained by printing the appropriate Value Assignment sheet posted on the Abbott Point of Care website at www.abbottpointofcare.com If any result is outside of the published range, the calibration verification should be repeated.  If the repeat calibration is still outside the acceptable range, troubleshooting is required. The lot number may not be used until the problem is resolved. A corrective action form will be filled out documenting actions taken for the resolution of the problem.  If the problem is not resolved, contact iSTAT technical service.

If calibration verification data are within the acceptance criteria, the AMR has been verified.

AMR VERIFICATION

College of American Pathologist accreditation requires the verification of the analytical measurement range. This requirement is non-applicable for coagulation assays.

Analytical Measureable Range (AMR) is 50-1000 seconds

Expected/ Reference Range

Expected Values

| Test/Abbreviation | Units | Reportable Range | Reference Range (PREWRM) | Reference Range (NONWRM) |
| --- | --- | --- | --- | --- |
| Activated Clotting Time/ACT+ | seconds | 50 - 1000* | 74 - 137 | 82- 152 |

LIMITATIONS OF PROCEDURE

The i-STAT Kaolin ACT test is to be used with fresh venous or arterial whole blood samples. The presence of exogenously added heparin, citrate, oxalate, or EDTA will interfere with test results. Poor technique in sample collection may also compromise the results. Samples drawn from insufficiently flushed catheters or from traumatic venipunctures may be contaminated with interfering substances. Samples should be collected into plastic syringes or tubes. Collection into glass may prematurely activate coagulation resulting in accelerated clotting times.

The analyzer should remain on a level surface with the display facing up during testing. If the analyzer is not level, the ACT result may be affected by more than 10%. A level surface includes running the handheld in the downloader/recharger.

Hemodilution may affect test results.

Platelet dysfunction, hereditary or acquired, may affect the results of this test. This includes the administration of pharmacological compounds known as platelet inhibitors which affect platelet function. Factor deficiencies, dysprothrombinemias, other coagulopathies, and other pharmacological compounds may also affect the results of this test.

The i-STAT ACT test is not affected by fibrinogen concentration in the range from 100 - 500 mg/dL, or sample temperature from 15 - 37ºC.

Interferences

An interferent is a substance which, if present at significant levels in the blood specimen being analyzed, will produce an error in the result of the analyte being measured.  There are no known interfering substances for ACT+ per manufacturer instructions.

MAINTENANCE

Routine maintenance other than cleaning is not required for the i-STAT System.  Clean and/or decontaminate as needed.

- Make sure the analyzer stays dry.  If liquids enter the electronics compartment, the battery compartment, or the cartridge port, the analyzer may be damaged.

- Clean the display screen with lintless tissue or alcohol pad when dirty.  Before use, squeeze the excess liquid from the pad.

- To decontaminate the i-STAT System, wipe the surface with Super Sani-Cloth disposable wipes or 10% bleach solution.

- Inspect and clean the cuvette opening as required.  Remove residual dried blood or other foreign matter using an alcohol wipe or a swab moistened with 10% bleach.  Wipe instrument with a water dampened cloth to remove bleach from the plastic surfaces.  Use Super Sani-Cloth disposable wipes to clean and disinfect areas contaminated with blood.  DO NOT use solvents or strong cleaning solutions as they may damage the instrument's plastic components.

- Exercise universal safety precautions at all times when handling the analyzer, cartridges, and peripherals to prevent exposure to blood-borne pathogens.

- The analyzer is NOT designed to be sterilized or autoclaved by any method.

- Dispose of cuvettes in biohazard trash.

UPDATING THE SOFWARE

The i-STAT 1 System is designed to eliminate operator influence on delivered results.  Due to continuous manufacturing process improvements, it is necessary to update standardization values from time to time to maintain long-term consistency of performance.  These updates are equivalent to manually adjusting calibration on a traditional laboratory analyzer.

New CLEW software is delivered twice a year to re-establish these standardization values and incorporate refinements to the internal quality  monitoring system.

New JAMS application software allows the i-STAT 1 Handheld to recognize any newly launched cartridge types and to perform any newly launched features.

For instructions on updating the i-STAT 1 Handheld, refer to the Technical Bulletin directly following section 18 in the System Manual.

TROUBLESHOOTING

When the analyzer detects a potential or real problem before the test cycle is initiated or at any time during the test cycle, a Quality Check Code number, the type of problem, and the step to be taken will be displayed.

Contact the Ancillary Testing Office at extension 53026 or 53034 for assistance.  Please provide the code number to assist in resolving the problem.  A loaner i-STAT Handheld will be available for use.

Alternative Procedure

Should the i-STAT System become inoperable for any reason, specimens should be collected and submitted to the laboratory for testing.

REFERENCES

i-STAT®1 System Manual, Abbott Point of Care, Revision Date: 07-Mar-13

I-Stat Procedure Manual for the I-Stat System, Revision Date:  22-Sep-16

i-STAT® Technical Bulletin Instructions for updating i-STAT®1Handheld Software, Art:

714547-00F Rev. Date 06/11/2008

i-STAT® Technical Bulletin Instructions for updating i-STAT®1Handheld Software, Art:

731336-00A Rev. Date 11-Mar-13

Cartridge and Test Information, Art:  715878-00N Rev. Date:  03/11/2013

ACT Kaolin, Art:  714183-00U rev. Date:  11-Mar-13

i-STAT® Technical Bulletin Calibration Verification and the iSTAT System, Art:  715209-00E Rev. Date:  06/11-08

i-STAT® Technical Bulletin Instructions for updating i-STAT®1Handheld Software, Art:

715878-00N Rev. Date 11-Mar-13

College of American Pathologists All Common Checklist

College of American Pathologists Point-of-Care-Testing Checklist