# POC.   CG8 Cartrige

Using the i-STAT® 1System

SCOPE

This procedure provides instruction on how to perform blood analysis using CG8+ iSTAT cartridge from patient blood specimens.

Complexity:

- This test is authorized for use at the Point of Care (POC), i.e., in patient care settings operating under a CLIA Certificate of Waiver, Certificate of Compliance, or Certificate of Accreditation.

Level of Personnel:

- All testing personnel who are qualified to performed nonwaived testing and who have successfully completed initial training and fulfilled the specific competency requirements for the complexity level of this test.

Testing Sites:

- Clinical sites approved and on file with the Pathology service’s POCT Division.

INTENDED USE

The i-STAT is a portable clinical analyzer used in conjunction with i-STAT cartridges to obtain definitive quantitative measurements of blood gas parameters utilizing i-STAT cartridges. The disposable cartridge contains all of the system’s components (reagents, sensors, and waste container).

PRINCIPLE

The i-STAT® Arterial Blood Gas (CG8) measurements are used in the diagnosis, monitoring, and treatment of respiratory disturbances and metabolic and respiratory-based acid-base disturbances as well as the  for analysis of sodium, potassium, glucose, ionized calcium, Hemoglobin and Hematocrit.

pH is measured by direct potentiometry. In the calculation of results for pH, concentration is related to potential through the Nernst equation.

PO2 is measured amperometrically. The oxygen sensor is similar to a conventional Clark electrode. Oxygen permeates through a gas permeable membrane from the blood sample into an internal electrolyte solution where it is reduced at the cathode. The oxygen reduction current is proportional to the dissolved oxygen concentration.

PCO2 is measured by direct potentiometry. In the calculation of results for PCO2, concentration

is related to potential through the Nernst equation.

Sodium, potassium and ionized calcium are measured by ion-selective electron potentiometry.  Potentiometry is the measurement of the difference in potential that exists between an indicator electrode and a reference electrode.  Activity is related to concentration through the activity coefficient.

Glucose is measured amperometrically.  Oxidation of glucose, catalyzed by the enzyme glucose oxidase, produces hydrogen peroxide (H2O2).  The liberated hydrogen peroxide is oxidized at the electrode to produce a current, which is proportional to the sample glucose concentration.

Hemoglobin is a protein substance found in red blood cells that transports oxygen, carbon dioxide throughout the body.  Hemoglobin is a calculated result.

Hemoglobin (g/dl) = hematocrit (decimal fraction) x34

Calculated Values

When a cartridge includes sensors for both pH and PCO2, bicarbonate (HCO3), total carbon dioxide (TCO2) and base excess (BE) are calculated.1

log HCO3 = pH + log PCO2 – 7.608

BEecf = HCO3 – 24.8 + 16.2 (pH – 7.4)

BEb = (1 - 0.014*Hb) * [ HCO3 - 24.8 + (1.43 * Hb + 7.7) * (pH - 7.4) ]

Anion Gap is calculated in the EC8+ and CHEM8+ cartridges as follows:

Anion Gap (EC8+) = (Na + K) – (Cl + HCO3)

Anion Gap (CHEM8+) = (Na + K) – (Cl + (TCO2 - 1)

The measured TCO2 test method is calibrated to the International Federation of Clinical Chemistry (IFCC) TCO2 reference method1 with an algorithm, based on the Henderson-Hasselbalch equation, which uses pH, PCO2, and ionic strength (Na) measurements.

Hemoglobin is a calculated result Hemoglobin (g/dl) = hematocrit (decimal fraction) x34

CLINICAL SIGNIFICANCE

pH is an index of the acidity or alkalinity of the blood with an arterial pH of <7.35 indicating an acidemia and >7.45 alkalemia.

PO2 (partial pressure of oxygen) is a measurement of the tension or pressure of oxygen dissolved in blood. Some causes for decreased values of PO2 include decreased pulmonary ventilation (e.g. airway obstruction or trauma to the brain), impaired gas exchange between alveolar air and pulmonary capillary blood (e.g. bronchitis, emphysema, or pulmonary edema), and alteration in the flow of blood within the heart or lungs (e.g. congenital defects in the heart or shunting of venous blood into the arterial system without oxygenation in the lungs).

PCO2 along with pH is used to assess acid-base balance. PCO2 (partial pressure of carbon dioxide), the respiratory component of acid-base balance, is a measure of the tension or pressure of carbon dioxide dissolved in the blood. PCO2 represents the balance between cellular production of CO2 and ventilator removal of CO2 and a change in PCO2 indicates an alteration in this balance. Causes of primary respiratory acidosis (increase in PCO2) are airway obstruction, sedatives and anesthetics, respiratory distress syndrome, and chronic obstructive pulmonary disease. Causes of primary respiratory alkalosis (decreased PCO2) are hypoxia (resulting in hyperventilation) due to chronic heart failure, edema and neurologic disorders, and mechanical hyperventilation.

HCO3 (bicarbonate), the most abundant buffer in the blood plasma, is an indicator of the buffering capacity of blood. Regulated primarily by the kidneys, HCO3 is the metabolic component of acid-base balance.Causes of primary metabolic acidosis (decrease in HCO3) are ketoacidosis, lactate acidosis (hypoxia), and diarrhea. Causes of primary metabolic alkalosis (increase in HCO3) are vomiting and antacid treatment.

Base excess of the extracellular fluid or standard base excess is defined as the concentration of titratable base minus the concentration of titratable acid when titrating the average intracellular fluid (plasma plus interstitial fluid) to an arterial plasma pH of 7.40 at PCO2 of 40 mmHg at 37°C. Excess concentration of base in the average ECF remains virtually constant during acute changes in the PCO2 and reflects only nonrespiratory component of pH-disturbances.

Anion gap is reported as the difference between the commonly measured cations sodium and potassium and the commonly measured anions chloride and bicarbonate. The size of the gap reflects unmeasured cations and anions and is therefore an analytical gap. Physiologically, a deficit of anions cannot exist. While relatively nonspecific, anion gap is useful for the detection of organic acidosis due to an increase in anions that are difficult to measure. Anion gap can be used to classify metabolic acidosis into high and normal anion gap types. Anion gap may be only slightly increased in diarrhea and renal failure, but elevated (often >25) due to an increase in organic anions in lactic acidosis, ketoacidosis (alcoholic, diabetic, starvation) and uremia, an increase in inorganic anions in uremia, and an increase in anions from drugs such a salicylate and carbenicillin or toxins such as methanol and ethanol.

TCO2 is a measure of carbon dioxide which exists in several states: CO2 in physical solution or loosely bound to proteins, bicarbonate (HCO3) or carbonate (CO3) anions, and carbonic acid (H2CO3). Measurement of TCO2 as part of an electrolyte profile is useful chiefly to evaluate HCO3 concentration. TCO2 and HCO3 are useful in the assessment of acid-base imbalance (along with pH and PCO2) and electrolyte imbalance.

Sodium measurements are used for monitoring electrolyte imbalances. Tests for sodium in the blood are important in the diagnosis and treatment of patients suffering from hypertension, renal failure or impairment, cardiac distress, disorientation, dehydration, nausea and diarrhea.

Potassium measurements are used in the diagnosis and monitoring of diseases and clinical conditions that manifest high and low potassium levels.  Tests for potassium in the blood are important in the diagnosis and treatment of patients suffering from hypertension, renal failure or impairment, cardiac distress, disorientation, dehydration, nausea and diarrhea.

Ionized calcium measurements are used in the diagnosis, monitoring, and treatment of conditions including, but not limited to, parathyroid disease, a variety of bone diseases, chronic renal disease, tetany, and disturbances related to surgical and intensive care.  Although most of the calcium in blood is bound to protein or complexed to smaller anionic species, the biologically active fraction of calcium is free ionized calcium. Through its role in a number of enzymatic reactions and in membrane transport mechanisms, ionized calcium is vitally important in blood coagulation, nerve conduction, neuromuscular transmission, and in muscle contraction.

Glucose measurements are used in the diagnosis, monitoring, and treatment of carbohydrate metabolism disorders including, but not limited to, diabetes mellitus, neonatal hypoglycemia, idiopathic hypoglycemia, and pancreatic islet cell carcinoma. Glucose is a primary energy source for the body and the only source of nutrients for brain tissue. Measurements for determination of blood glucose levels are important in the diagnosis and treatment of patients suffering from diabetes and hypoglycemia.

Hematocrit measurements can aid in the determination and monitoring of normal or abnormal total red cell volume status including, but not limited to, conditions such as anemia, erythrocytosis, and blood loss related to trauma and surgery.  Hematocrit is a measurement of the fractional volume of red blood cells. This is a key indicator of the body’s state of hydration, anemia or severe blood loss, as well as the blood’s ability to transport oxygen.

Hemoglobin:  Decreased hemoglobin is associated with blood loss, hemolytic anemia, and bone marrow suppression.  Elevated hemoglobin levels are associated with dehydration, polycythemia, congestive heart failure and severe burns.

SPECIMEN COLLECTION AND HANDLING

Use standard biohazard precautions when collecting any blood specimen.  Wear gloves during specimen collection and while testing.  Dispose of needles in appropriate puncture resistant containers.

The I-Stat CG8+ cartridge test can be performed using venous or arterial samples.

Specimen Labeling

The patient is identified by the patient ID bracelet, by asking patient to state their name and SSN, or by identification from a family member or nurse in lieu of asking the patient; if he is unresponsive. Because the specimen is analyzed immediately after collection; tested in the same room where the patient procedure is being performed; tested by one authorized user, and then discarded, there are no requirements for specimen labeling.  In the event the specimen is taken into another procedure room or testing performed by a second authorized user, the specimen container must be labeled with patient’s full name, full SSN, source, date and time of collection and collector’s initials.

COLLECTION

If a syringe is used, it should have a 23 gauge needle or larger.  DO NO use excessive force when expelling the blood specimen through the needle.  This may cause hemolysis.

Arterial puncture:  Plain syringe or blood gas syringe with heparin labeled for the assays performed or with the least amount of heparin to prevent clotting (10 IU heparin/mL of blood)

Blood Samples must be collected according to the following procedures to assure the integrity of the fresh whole blood sample.

Syringe sample, from indwelling line:

NOTE: The amount of blood required to adequately flush the line until it is free of contaminants is dependent on the amount of solution contained within the line. A typical heparin lock will require approximately 5 cc to clear the line. Greater volumes will be required to clear longer lines.

- In-Dwelling Line

- Back flush line with sufficient amount of blood to remove intravenous solution, heparin, or medications that may contaminate the sample. Recommendation: three to six times the volume of the catheter, connectors, and needle.

- Arterial/Venous Specimens

For cartridge testing of blood gases, fill a plain syringe or fill a blood gas syringe, labeled for the assays to be performed to the recommended capacity, or use the least amount of liquid heparin anticoagulant that will prevent clotting. Under filling syringes containing liquid heparin will decrease results due to dilution and will decrease ionized calcium results due to binding.

Mix blood and anticoagulant by rolling syringe between palms for at least 5 seconds and then inverting the syringe repeatedly for at least 5 seconds. For blood gas testing, avoid or remove immediately any air drawn into syringe to maintain anaerobic conditions.

Suitable Specimens for CG8+:

- Only fresh whole blood collected in a plastic syringe without anticoagulant. Test immediately.

- Fill syringes for correct blood-to-heparin ratio.  Test within 10 minutes of collection.

- Sample requirement is 40 µl of fresh whole blood (venous or arterial) without clot activators, or serum separators collected in a plastic syringe properly

- The collection device cannot contain clot activators or serum separators.

- The sample should be immediately dispensed into the sample well of a test cartridge.

- If a repeat measurement is needed, sample can be retested.  If 10 minutes or more have elapsed, a fresh sample must be obtained.

Criteria For Specimen Rejection

- Evidence of clotting

- Specimens collected in vacuum tubes with anticoagulant other than lithium or sodium heparin

- Syringe for pH, PCO2, and PO2 with air bubbles in sample

- Incompletely filled vacuum tube for the measurement of ionized calcium or PCO2

- Other sample types such as urine, CSF, and pleural fluid

Precautions: Avoid the Following Circumstances

- Drawing a specimen from an arm with an I.V.

- Stasis (tourniquet left on longer than one minute before venipuncture)

- Extra muscle activity (fist pumping)

- Hemolysis (alcohol left over puncture site, or a traumatic draw)

- Icing before filling cartridge

- Time delays before filling cartridge

- Exposing the sample to air when measuring pH, PCO2, and PO2

REAGENTS - SUPPLIES - EQUIPMENT

i-STAT® 1 Analyzer

i-STAT® CG8+ Cartridge.  Store cartridges at 2 to 8°C (35-46°F)

Bulk storage in Pathology, room D120 refrigerator

Cartridges are sealed in individual pouches.  Store the main supply of cartridges at a temperature between 2 to 8°C (35 to 46°F). Do not allow cartridges to freeze.   Cartridges may be stored at room temperature (18 to 30°C or 64 to 86°F) for 2 months.   When cartridges are removed from the refrigerator, write the room temperature expiration date on the cartridge/box which is 2 months from opened date. Cartridges should not be returned to the refrigerator once they have been at room temperature, and should not be exposed to temperatures above 30°C (86°F).   If the pouch has been punctured, the cartridge should not be used. Cartridges should remain in pouches until time of use.  Do not use after the labeled expiration date.

i-STAT® TriControls Level 1 and Level 3

Store at 2 to 8°C (35-46°F)

i-STAT® TriControls Calibration Verification Kit

Store at 2 to 8°C (35-46°F)

Contents

Each cartridge contains a buffered aqueous calibrant containing know concentrations of each analyte; one reference electrode; one ion-selective electrode each for sodium, potassium, and ionized calcium, one amperometric electrode with immobilized glucose oxidase for glucose, one electrode for direct potentiometry, and one conductometric sensor for hematocrit.

A whole blood sample of approximately 1 to 3 drops is dispensed into the cartridge sample well and the sample well is sealed.

NOTE:  All new lots of cartridges will be validated before

patient testing is initiated on that lot.

QUALITY CONTROL

- Liquid quality control is performed monthly on each analyzer using i-STAT TriControls Level 1 and Level 3.   Quality control is due on the first Wednesday of every month. All testing personnel must perform quality control on a rotating basis. If quality control is not performed, the instrument will lock-out users until QC is performed.

i-STAT TriControls

Controls are stored refrigerated at 2 to 8°C (35° to 46°F) until expiration date on the box.  Do not use after expiration date on the box and ampules.

Procedure for testing cartridges with I-STAT Controls

- Prior to use, allow one vial each of Level 1 and Level 3 controls to stand at room temperature for a minimum of 30 minutes.

- Mix the control vial for 10 seconds or 10 shakes.

- Tap the top of the vial to dislodge any fluid in the top of the vial.

- Wearing gloves and using gauze, snap the top of the vial.

- Use a plastic pipette, syringe, or capillary tube without anticoagulant to transfer the solution to the cartridge. DO NOT REMOVE AIR BUBBLES FROM SYRINGE.

- Dispense solution in the cartridge.

- Immediately seal the cartridge and insert it into a handheld.

- Results must fall within the expected ranges published in the Value Assignment Sheets.  The i-STAT analyzer is programmed to automatically determine if the results are within the published quality control ranges and flag quality control results as PASS or FAIL.  If results are within the expected ranges, use the cartridges as needed.  If quality control results FAIL, contact the ATC.

Target values are printed on a Value Assignment Sheet posted on the APOC website at www.abbottpointofcare.com.  Ensure that the lot number printed on the Value Assignment Sheet matches the lot number on the label of the ampule and that the software version above the target value table matches the software version in the handheld.

The ranges displayed represent the maximum deviation expected when controls and cartridges are performing properly.  Should results outside these ranges be obtained, refer to the Troubleshooting section that follows the Procedure for Testing Controls in the System Manual. If the repeated QC is still outside the acceptable range, the analyzer may not be used until the problem is resolved. A corrective action form will be filled out documenting actions taken for the resolution of the problem.  If the problem is not resolved, contact iSTAT technical service.

Performing Internal Electronic Simulator Test

The performance of each handheld analyzer in the i-STAT System is automatically verified using an internal Electronic Simulator every 8 hours of use.  The Electronic Simulator simulates two levels of electrical signals that stress the analyzer’s cartridge signal detection function both below and above measurement ranges.

The internal Electronic Simulator test cycle is automatically activated when a cartridge is inserted after the customized interval is reached.  If the analyzer passes the simulator test, the cartridge test cycle proceeds.  If not, the analyzer displays “ELECTRONIC SIMULATOR FAIL”.    The same cartridge can be re-inserted immediately after the FAIL message is displayed.  If the analyzer fails the simulator test again, contact the Ancillary Testing Office at extension 53026 or 53034 for assistance.

Performing External Electronic Simulator Test

This process will be completed by the Ancillary Testing Coordinator or designee.  The External Electronic Simulator and i-STAT® 1 System Manual will be stored in the Ancillary Testing Office.

Checking Thermal Probes

This process will be completed by the Ancillary Testing Coordinator or designee and by using the instructions in the i-STAT® 1 System Manual.  Thermal Probes will be checked every six months.

Performing Internal QC

The i-STAT has a lock-out in place to automatically run the internal electronic QC every 8 hours of patient testing. The analyzer will not proceed with patient testing if the internal QC does not pass. Internal QC is stored in RALS. Any QC failure and corrective action taken is recorded on the corrective action form.

###### Proficiency Testing

Proficiency testing will be performed 3 times per year using samples obtained from the College of American Pathologists.  Testing will be performed by personnel who routinely perform patient testing.  Comparison of results with any other testing site prior to submission to the CAP and referral of PT specimens to outside institutions is prohibited.

Proficiency testing will be used as a system to compare the accuracy of the testing procedure and the proficiency of the users with other laboratories.  All iSTAT instruments will be tested with the CAP survey material and results reported.

All authorized operators regularly using the iSTAT instrument will complete testing of the CAP survey.  Survey samples will be rotated among staff as evenly as possible with consideration to patient workload at the time.

CALIBRATION

Calibration of the i-STAT® system is automatically performed as part of the test cycle on each cartridge and does not require any action by the operator.

AMR VERIFICATION

Joint Commission requires the verification of the analytical measurement range.

i-STAT TriControls Calibration Verification is a 5-point linearity matrix that is used to verify instrument linearity at least semiannually.

CALIBRATION/VERIFICATION/LINEARITY

VERIFICATION/CORRELATION/LINEARITY OF INSTRUMENT/METHOD ASSESSMENT

- When more than one instrument/method is used to test a given analyte, correlation will be performed against each instrument/method twice a year.

- CAP surveys or patient samples will be used for comparison.

- Instrument verifications are performed twice a year using matrix appropriate materials.

- Calibration/verification/linearity samples are tested in duplicate and acceptability criteria are defined for each analyte.

- When results fail, a corrective action form is filled out.  Affected lot numbers cannot be used for patient testing until errors are resolved.

- If repeat calibration/verification fails, contact technical support.

Calibration/verification/linearity is a procedure intended to validate instrument performance according to manufacturer expectations. Using the i-STAT System for nonwaived testing and following manufacturer’s recommendations, calibration verification will take place for the following:

- Before any new i-STAT System is placed into use.

- At changes of reagent lots upon initial use of the i-STAT System until lot-to-lot variation is proven to be consistent.  Verify the new lot using the Calibration Verification Kit for the first six lot numbers comparing each to the previous lot(s).  Lot to lot variation must be within the evaluation limit of +0.3 mg/dL or 15%, whichever is greater.

- To verify that results have not been affected by major maintenance or repair.

- To troubleshoot when control values are out-of-range or reflect an unusual trend or shift.

- As recommended by manufacturer

- At least every six months

Calibration/linearity is performed semiannually by testing all five levels of i-STAT TriControls Calibration Verification material in duplicate.   Target values are obtained by printing the appropriate Value Assignment sheet posted on the Abbott Point of Care website at www.abbottpointofcare.com If any result is outside of the published range, the calibration verification should be repeated.  If the repeat calibration is still outside the acceptable range, troubleshooting is required. The lot number may not be used until the problem is resolved. A corrective action form will be filled out documenting actions taken for the resolution of the problem.  If the problem is not resolved, contact iSTAT technical service.

If calibration/verification/linearity data are within the acceptance criteria, the AMR has been verified.

PROCEDURE

Start Up

PATIENT TESTING

Specimens will be analyzed immediately after collection.  Only fresh whole blood collected in plastic syringes or tubes with anticoagulant may be used for the CG8+ test. If samples are collected in plastic tubes, use a plastic syringe to transfer sample to the testing device. Testing must be performed on the i-STAT System with no delays

Cartridge Handling

- Do not remove the cartridge from its protective pouch until it is at room temperature (18-30°C or 64-86° F).

- Use a cartridge immediately after removing it from its protective pouch.  Prolonged exposure may cause a cartridge to fail a Quality Check.

- Do not contaminate the contact pads with fingerprints.

- Do not apply pressure to the central area of the label as the calibrant pack inside could burst prematurely.

- Do not block the air vent as the sample will not flow to the fill mark and the calibrant solution will not flow to the sensors.

- Do not use a cartridge on which blood or any other fluid has spilled.  Avoid filling cartridges on surfaces that may cause the cartridge to pick up fibers, fluid, or debris that may lodge in the analyzer.

Filling and Sealing Cartridge

- Place cartridge on a flat surface.

- Direct the tip of the dispensing device into the sample well.

- Dispense sample slowly and steadily until it reaches the fill mark indicated on the cartridge label.  Leave some sample in the sample well.

- Press the rounded end of the closure until it snaps into place.  Slightly lift finger or thumb and ensure that the cartridge is closed before completely removing finger or thumb from the closure.

Inserting and Removing the Cartridge from the Analyzer

- Inserting

  - Align the cartridge with the contact pads facing up and toward the cartridge port

  - Push the cartridge slowly and smoothly into the cartridge port until it clicks into place.

- Removing

  - Do not attempt to remove the cartridge while the message “Cartridge Locked” remains on the screen.

  - When the results are displayed, pull the cartridge straight out of the analyzer.

- Dispose of the cartridge in a container for biohazards.

Performing Patient Analysis

- Press the On/Off button

- Press 2, for i-STAT Cartridge

- Press SCAN and hold until the operator barcode is read, or manually enter your operator ID

- Press SCAN to enter the patient ID from the wrist band, or manually enter the patient ID

- Repeat patient ID entry to verify correct entry, if necessary.

- Press SCAN and hold to enter the cartridge lot number

- Unwrap the cartridge, touching only the sides of the cartridge

- Fill the cartridge according to instructions above in “Filling and Sealing Cartridge”.

- Push the sealed cartridge into the handheld port until it clicks into place. Ensure that analyzer is kept on a flat surface during testing. Do not attempt to remove the cartridge while the message “Cartridge Locked” remains on the screen.

- Wait for the test to complete.

- Review results

- Follow the screen prompt and remove the cartridge.  Discard the cartridge in a bio-hazard bag.

- Place i-Stat into downloader/recharger to transmit data to VISTA/CPRS.

Shutdown

Press to turn off Handheld.

REPORTING OF RESULTS

Results are automatically reported in CPRS (with appropriate reference ranges when applicable) when the analyzer is returned to its base for downloading. The operator performing the test is responsible for proper evaluation, documentation of results, and treatment necessary for the patient.  The results will be entered in the lab report section of the patient electronic medical record (EMR) via the RALS-Web3 interface.

Reference Ranges are listed below:

Reference Ranges

| Test | Reference Ranges | Reportable Ranges Analytical Measurable Range (AMR) | Critical Ranges |
| --- | --- | --- | --- |
|  |  |  | Low High |
| Sodium | 136.0 - 145.0 mmol/L | 100 – 180 mmol/L | < 120.0 mmol/L - >160.0 mmol/L |
| Potassium | 3.50 - 5.00 mmol/L | 2.0 - 9.0 mmol/L | < 3.0 mmol/L - > 6.0 mmol/L |
| Ionized Calcium | 1.090 - 1.300 mmol/L | 0.25 – 2.50 mmol/L | < 0.9 mmol/L - > 1.5 mmol/L |
| Glucose | 72.0 – 99.0 mg/dL | 20 – 700 mg/dL | < 45 mg/dL - > 500 mg/dL |
| pH | 7.350 – 7.450 | 6.50 – 8.20 | < 7.250 - > 7.600 |
| pCO2 | 35.00 – 45.00 mmHg | 5 – 800 mmHg | < 25.00 mmHg - > 60.00mmHg |
| pO2 | 80.0 – 100.0 mmHg | 5 – 800 mmHg | < 50.0 mmHg - N/A |
| TCO2 | 24.0 – 30.0 mmol/L | 5 – 50 mmol/L | < 10.0 mmol/L - > 42.0mmol/L |
| HCO3 | 20.0 – 26.0 mmol/L | 20.0 – 26.0 mmol/L | N/A |
| BE |  | -30 - +30 | N/A |
| sO2 | 85% – 100% | N/A | N/A |
| Hemoglobin | 13.10 – 16.80 g/dL | 3.4 – 25.5 g/dL | < 7.0 g/dL - N/A |
| Hematocrit | 38.2 – 48.4 %PCV | 10 - 75 %PCV | N/A |

INTERPRETATION

Critical Results

No critical note documention is required due to results given directly to phyisician during procedure.

INTERFERENCES AND LIMITATIONS OF THE TEST

The i-STAT CG8+ test is to be used with fresh venous or arterial whole blood samples. Poor technique in sample collection may also compromise the results. Samples drawn from insufficiently flushed catheters or from traumatic venipunctures may be contaminated with interfering substances. Samples should be collected into plastic syringes or tubes.

- EDTA or Sodium heparin may not be used for any sample run on the i-STAT.

- Sodium heparin may increase sodium results up to 1 mEq/L.

- Hemodilution of the plasma by more than 20% associated with priming cardiopulmonary bypass pumps or other fluid administration therapies using certain solutions may cause clinically significant error on sodium, chloride, ionized calcium and pH results.  These errors are associated with solutions that do not match the ionic characteristics of plasma.  To avoid these errors when hemodiluting by more than 20%, use physiologically balanced multi-electrolyte solutions containing low-mobility anions (e.g. gluconate) rather than solutions such as normal saline or ringer’s lactate.

- If heparinized whole blood is allowed to stand before testing, potassium values will first decrease slightly, then increase over time.  Potassium values will increase in iced specimens.

- Potassium values from serum samples can be 0.1 to 0.7 mEq/L higher than plasma/whole blood samples due to potassium released from platelets and red blood cells during the clotting process.

- The dependence of the i-STAT glucose with respect to pH is as follows:  pH values below 7.4 at 37° C decrease glucose results by approximately 0.9 mg/dL per 0.1 pH units.  Values above 7.4 at 37° C increase results by approximately 0.8 mg/dL per 0.1 pH unit.

- The dependence of the i-STAT glucose with respect to pO2 is as follows:  Oxygen levels of less than 20 mmHg at 37°C may decrease results.

- Mean Cell Volumes of K3EDTA anticoagulated blood are approximately 2-4% less than K2EDTA anticoagulated blood.

- The measurement of certain blood samples with high erythrocyte sedimentation rates (ESR) may be affected by analyzer angle.  While testing blood samples, beginning ninety seconds after the cartridge is inserted, the analyzer should remain level until a result is obtained. A level surface includes running the handheld in the downloader/recharger.

- Grossly elevated white blood cell counts may increase results for hematocrit.

- Hematocrit results are affected by the level of total protein as follows:

| Displayed Result | TP < 6.5 g/dL | TP > 8.0 g/dL |
| --- | --- | --- |
| Hct < 40% | Hct decreased by ~1% for each decrease of 1 g/dL TP | Hct increased by ~1% for each increase 1 g/dL TP |
| Hct >40% | Hct decreased by ~0.75% for each decrease of 1 g/dL TP | Hct increased by ~0.75% for each increase 1 g/dL TP |

Total protein levels may be low in neonatal and burn patient populations, as well as in additional clinical populations listed in Statland.3 Total protein levels may also be decreased in patients undergoing cardiopulmonary bypass (CPB) or ECMO, and with patients receiving large volumes of saline-based IV fluids. Care should be taken when using hematocrit results from patients with total protein levels below the adult reference range (6.5 to 8 g/dL).

The CPB sample type can be used to correct the hematocrit result for the dilutional effect of the pump prime in cardiovascular surgery. The CPB algorithm assumes that cells and plasma are diluted equally and that the pump priming solution has no added albumin or other colloid or packed red blood cells. Since perfusion practices vary, it is recommended that each practice verify the use of the CPB sample type and the length of time in which the CPB sample type should be used during the recovery period. Note that for hematocrit values above 30 %PCV, the CPB correction is ≤1.5 %PCV; the size of the correction at this level should not impact transfusion decisions.

- Abnormally high lipids may increase results for hematocrit.  Interference from lipids will be about two thirds the size of the interference from protein.

- The sample electrolyte concentration is used to correct the measured conductivity prior to reporting hematocrit results.  Factors that affect sodium will therefore also affect hematocrit.

- Bromide (37.5 mmol/L) is known to result in an increased rate of star (***) outs.  Bromide has been tested at two levels: the CLSI recommended level and a therapeutic plasma concentration level of 2.5 mmol/L. The latter is the peak plasma concentration associated with halothane anesthesia, in which bromide is released. APOC has not identified a therapeutic condition that would lead to levels consistent with the CLSI recommended level. Bromide at a concentration of 37.5 mmol/L increased the rate of HCT star (***) outs, while a therapeutic range of bromide (2.5 mmol/L) did not significantly interfere with i-STAT HCT results.

- Low hematocrit results can be caused by contamination of flush solutions in an arterial or venous line.

- Venous stasis (prolonged tourniquet application) and forearm exercise may increase ionized calcium due to a decrease in pH caused by localized production of lactic acid.   Exposing the sample to air will cause an increase in pH due to the loss of CO2 which will decrease ionized calcium.

- Heparin binds calcium.  Each unit of heparin added per mL of blood will decrease ionized calcium by 0.01 mmol/L.  Therefore, the correct ratio of heparin anticoagulant to blood must be achieved during sample collection.  Intravenous injection of 10,000 units of heparin has been shown in adults to cause a significant decrease of ionized calcium of about 0.03 mmol/L.  Under-filling syringes will cause higher heparin-to-blood ratios which will decrease ionized calcium results due to binding.

- Exposure of the sample to air will cause an increase in pO2 when values are below 150 mmHg and a decrease in pO2 when values are above 150 mmHg (approximate pO2 of room air.)

- Standing anaerobically at room temperature will decrease pH at a rate of 0.03 per hour, will increase pCO2 by approximately 4 mmHg per hour and will decrease pO2 at a rate of 2 – 6 mmHg per hour.5 This will cause HCO3 and TCO2 to be over-estimated, due to metabolic processes.

- Do not use a cold cartridge – pO2 results may be falsely decreased if the cartridge is cold.

- Do not ice samples before testing – PO2 results may be falsely elevated in cold samples.

- sO2 values calculated from a measured PO2 and an assumed oxyhemoglobin dissociation curve may differ significantly from the direct measurement.

- Exposing the sample to air allows CO2 to escape which causes pCO2 to decrease and pH to increase and HCO3 and tCO2 to be under-estimated.  The use of partial-draw tubes (evacuated tubes that are adjusted to draw less than the tube volume e.g., a 5 cc tube with enough vacuum to draw only 3 cc)) is not recommended for use with the i-STAT because of the potential for decreased measured pCO2 results and calculated HCO3 and TCO2 values.  Under-filling blood collection tubes may also cause decrease pCO2 results.  Care must also be taken to eliminate “bubbling” of the sample with a pipette when filling a cartridge to avoid the loss of CO2 in the blood.

- Samples drawn from insufficiently flushed catheters or from traumatic venipunctures may be contaminated with interfering substances. Avoid hemolysis by allowing residual alcohol to dry over the puncture site; and by discarding a sample from a traumatic draw.

- Samples should be collected into plastic syringes or tubes.  Collection into glass may prematurely activate coagulation resulting in accelerated clotting times. The analyzer should remain on a level surface with the display facing up during testing.

- The following substances are known to interfere with the i-STAT sodium assay:

| Substance | Test Concentration mmol/L | Interference |
| --- | --- | --- |
| Bromide | 37.5 | Increased i-STAT Sodium results. See Note below. |

NOTE 1:  Bromide has been tested at two levels; the CLSI recommended level and a concentration of 2.5 mmol/L.  The latter is the peak plasma concentration associated with halothane anesthesia, in which bromide is released.  APOC has not identified a therapeutic condition that would lead to levels consistent with the CLSI recommended level.  Bromide at a concentration of 37.5 mmol/L increased i-STAT sodium results, while a therapeutic range of bromide (2.5 mmol/L) did not significantly interfere with i-STAT sodium results.

- The following substances are known to interfere with the i-STAT potassium assay:

| Substance | Test Concentration (mmol/L) | Interference |
| --- | --- | --- |
| Bromide | 37.5 | Increased rate of star (***) outs. See Note below. |

Note: Bromide has been tested at two levels: the CLSI recommended level and a therapeutic plasma concentration level of 2.5 mmol/L. The latter is the peak plasma concentration associated with halothane anesthesia, in which bromide is released. APOC has not identified a therapeutic condition that would lead to levels consistent with the CLSI recommended level. Bromide at a concentration of 37.5 mmol/L increased the rate of potassium star (***) outs, while a therapeutic range of bromide (2.5 mmol/L) did not significantly interfere with i-STAT potassium results.

- The following substances are known to interfere with the i-STAT glucose assay:

| Substance | Test Concentration mmol/L | Interference |
| --- | --- | --- |
| Acetaminophen | 1.32 | Increased i-STAT Glucose results. See Note below. |
| Acetylcysteine | 10.2 | Decreased i-STAT Glucose results. See note below. |
| Bromide | 37.5 | Decreased i-STAT Glucose results. See Note below. |
| Bromide (therapeutic) | 2.5 | Decreased i-STAT Glucose results. See Note below. |
| Hydroxyurea | 0.92 | Increased i-STAT Glucose results. Use another method. |
| Thiocyanate | 6.9 | Decreased i-STAT Glucose results by approximately 7 mg/dL |

Note 1:  Acetaminophen has been shown to interfere with glucose results in the i-STAT 6+, EC8+, EC4+ and G products, at a concentration prescribed by the CLSI guideline, 1.32 mmol/L, which represents a toxic concentration of acetaminophen. Acetaminophen at 0.132 mmol/L, which represents the upper end of the therapeutic concentration, has been shown not to significantly interfere with i-STAT glucose results for all i-STAT cartridges. Acetaminophen at a test level of 1.32 mmol/L (toxic concentration) has been shown not to significantly interfere with glucose results in the i-STAT CHEM8+ and CG8+ products.

NOTE 2:  Acetylcysteine has been tested at two levels; the CLSI recommended level and a concentration of 0.30 mmol/L. The latter is 3 times the peak plasma therapeutic concentration associated with treatment to reverse acetaminophen poisoning.  APOC has not identified a therapeutic condition that would lead to levels consistent with the CLSI recommended level.  Acetylcysteine at a concentration of 10.2 mmol/L decreased i-STAT glucose results, while acetylcysteine at a concentration of 0.3 mmol/L did not significantly interfere with i-STAT glucose results.

NOTE 3:  Bromide has been tested at two levels; the CLSI recommended level and a therapeutic plasma concentration of 2.5 mmol/L.  The latter is the peak plasma concentration associated with halothane anesthesia, in which bromide is released.  APOC has not identified a therapeutic condition that would lead to levels consistent with the CLSI recommended level.  Bromide at a concentration of 37.5 mmol/L significantly decreased i-STAT glucose results.  Bromide tested at concentrations of 2.5 and 37.5 mmol/L decreased i-STAT glucose results.

NOTE 4:  Hydroxyurea is a DNA synthesis inhibitor used in the treatment of various forms of cancer, sickle cell anemia, and HIV infection.  This drug is used to treat malignancies including melanoma, metastatic ovarian cancer, and chronic myelogenous leukemia.  It is also used in the treatment of polycythemia vera, thrombocytopenia, and psoriasis.  At typical doses ranging from 500 mg to 2 g/day, concentrations of hydroxyurea in patients’ blood may be sustained at approximately 100 to 500 µmol/L.  Higher concentrations may be observed soon after dosing or at higher therapeutic doses.

Note 5:  The dependence of the i-STAT glucose with respect to pH is as follows: Values below 7.4 at 37°C decrease results by approximately 0.9 mg/dL (0.05 mmol/L) per 0.1 pH units. Values above 7.4 at 37°C increase results by approximately 0.8 mg/dL (0.04 mmol/L) per 0.1 pH units.

Note 6:  The dependence of the i-STAT glucose with respect to PO2 is as follows: Oxygen levels of less than 20 mmHg (2.66 kPa) at 37°C may decrease results.

- The following substances are known to interfere with the i-STAT Ionized Calcium

| Substance | Test Concentration mmol/L | Interference |
| --- | --- | --- |
| Acetaminophen | 1.32 | Decreased i-STAT Ionized Calcium results. See Note below. |
| Acetylcysteine | 10.2 | Decreased i-STAT Ionized Calcium results. See Note below. |
| Bromide | 37.5 | Increased i-STAT Ionized Calcium results. See Note below. |
| Magnesium | 1.0 | Increased i-STAT Ionized Calcium results by approximately 0.04 mmol/L. |
| Lactate | 6.6 | Decreased i-STAT Ionized Calcium results by approximately 0.07 mmol/L. |
| Salicylate | 4.34 | Decreased i-STAT Ionized Calcium results. See Note below. |
| Salicylate (therapeutic) | 0.5 | Decreased i-STAT Ionized Calcium results by approximately 0.03 mmol/L. See Note below. |
| Thiocyanate | 6.9 | Decreased i-STAT Ionized Calcium results. Use another method. |

Note 1: Acetaminophen has been shown to interfere at a concentration proscribed by CLSI guideline, 1.32 mmol/L, which represents a toxic concentration.  Acetaminophen at 0.132 mmol/L, which represents the upper end of the therapeutic concentration, has been shown not to significantly interfere with i-STAT ionized calcium results.

Note 2:  Acetylcysteine has been tested at two levels; the CLSI recommended level and a concentration of 0.30 mmol/L.  The latter is 3 times the peak plasma therapeutic concentration associated with treatment to reverse acetaminophen poisoning.  APOC has not identified a therapeutic condition that would lead to levels consistent with the CLSI recommended level.  Acetylcysteine at a concentration of 10.2 mmol/L decreased i-STAT ionized calcium results, while acetylcysteine at a concentration of 0.3 mmol/L did not significantly interfere with i-STAT ionized calcium results.

Note 3:  Bromide has been tested at two levels; the CLSI recommended level and a concentration of 2.5 mmol/L.  The latter is the peak plasma concentration associated with halothane anesthesia, in which bromide is released.  APOC has not identified a therapeutic condition that would lead to levels consistent with the CLSI recommended level.  Bromide at a concentration of 37.5 mmol/L increased i-STAT ionized calcium results while bromide at a concentration of 2.5 mmol/L did not significantly interfere with i-STAT ionized calcium results.

Note 4:  Salicylate has been shown to significantly decrease ionized calcium results at a concentration proscribed by the CLSI guideline, 4.34 mmol/L, which represents a toxic concentration.  Salicylate at 0.5 mmol/L, which represents the upper end of the therapeutic concentration, has been shown to decrease ionized calcium results by approximately 0.03 mmol/L.

Suppressed Results

There are three conditions under which the i-STAT System will not display results:

- Results outside the System’s reportable ranges are flagged with a < or >, indicating that the result is below the lower limit or above the upper limit of the reportable range respectively.  (See the table of Reportable Ranges.)  The < > flag indicates that the results for this test were dependent on the result of a test flagged as either > or <.

Action:

Analyze the specimen again using another cartridge or send specimen(s) to the laboratory for analysis, if necessary.

- If stars (***) are displayed instead of a result, it means that a test failed internal quality checks.

Action:

Analyze the specimen again using another cartridge. If the result is suppressed again, send specimen(s) to the laboratory for analysis.

- A Quality Check message will be reported instead of results if the handheld detects a problem with the sample, calibrant solution, sensors, or mechanical or electrical functions of the handheld during the test cycle.

Action:

Take the action displayed with the message that identifies the problem. Contact the Ancillary Testing Coordinator for assistance.

TRANSMITTING INFORMATION

- All meters should be returned to the base unit upon completion of quality control and patient testing.

- Place the meter back in the base unit and verify that the screen displays “Communication in Progress” message.  This verifies patient results are transmitting to the computer network server and information from the RALSWeb3 database is being returned to the meter.

- Do not move analyzer while the message “Communication in Progress” is displayed.

- Results will transmit to the data management system.  Results not needing review will automatically download to the patient’s medical record.  Results needing review will be marked as flagged and will not download until reviewed by the Ancillary Testing Coordinator.

MAINTENANCE

Routine maintenance other than cleaning is not required for the i-STAT System.  Clean and/or decontaminate as needed.

- Make sure the analyzer stays dry.  If liquids enter the electronics compartment, the battery compartment, or the cartridge port, the analyzer may be damaged.

- Clean the display screen with lintless tissue or alcohol pad when dirty.  Before use, squeeze the excess liquid from the pad.

- To decontaminate the i-STAT System, wipe the surface with Super Sani-Cloth disposable wipes or 10% bleach solution.

- Inspect and clean the cuvette opening as required.  Remove residual dried blood or other foreign matter using an alcohol wipe or a swab moistened with 10% bleach.  Wipe instrument with a water dampened cloth to remove bleach from the plastic surfaces.  Use Super Sani-Cloth disposable wipes to clean and disinfect areas contaminated with blood.  DO NOT use solvents or strong cleaning solutions as they may damage the instrument's plastic components.

- Exercise universal safety precautions at all times when handling the analyzer, cartridges, and peripherals to prevent exposure to blood-borne pathogens.

- The analyzer is NOT designed to be sterilized or autoclaved by any method.

- Dispose of cuvettes in biohazard trash.

UPDATING THE SOFWARE

The i-STAT 1 System is designed to eliminate operator influence on delivered results.  Due to continuous manufacturing process improvements, it is necessary to update standardization values from time to time to maintain long-term consistency of performance.  These updates are equivalent to manually adjusting calibration on a traditional laboratory analyzer.

New CLEW software is delivered twice a year to re-establish these standardization values and incorporate refinements to the internal quality  monitoring system.

New JAMS application software allows the i-STAT 1 Handheld to recognize any newly launched cartridge types and to perform any newly launched features.

For instructions on updating the i-STAT 1 Handheld, refer to the Technical Bulletin in the System Manual.

TROUBLESHOOTING

When the analyzer detects a potential or real problem before the test cycle is initiated or at any time during the test cycle, a Quality Check Code number, the type of problem, and the step to be taken will be displayed.

Contact the Ancillary Testing Office at extension 53026 or 53034 for assistance.  Please provide the code number to assist in resolving the problem.  A loaner i-STAT Handheld will be available for use.

Alternative Procedure

Should the i-STAT System become inoperable for any reason, specimens should be collected and submitted to the laboratory for testing.

PROCEDURAL NOTES  NA

RELATED DOCUMENTS  NA

REFERENCES

i-STAT®1 System Manual, Abbott Point of Care.

I-Stat Procedure Manual for the I-Stat System.

i-STAT® Technical Bulletin Instructions for updating i-STAT®1Handheld Software, Art:

714547-00F

i-STAT® Technical Bulletin Instructions for updating i-STAT®1Handheld Software, Art:

731336-00A

Cartridge and Test Information, Art:  715878-00N

ACT Kaolin, Art:  714183-00U

i-STAT® Technical Bulletin Calibration Verification and the iSTAT System, Art:  715209-00E Rev.

i-STAT® Technical Bulletin Instructions for updating i-STAT®1Handheld Software, Art:

715878-00N

McPherson MD MSc, R; Pincus MD PhD, M. 2017. Henry’s Clinical Diagnosis and Management of by Laboratory Methods 23rd Edition. Pages 481-508. St Louis, MO: Elsevier Inc.

Digital Copy link: https://www.clinicalkey.com/#!/browse/book/3-s2.0-C20130143425

College of American Pathologists All Common Checklist

College of American Pathologists Point-of-Care-Testing Checklist