Abbott i-STAT Operator Training and Annual Competency Checklist

| Name: | Date: |
| --- | --- |
| ID #: | Assigned Location: Respiratory Therapy |

This individual has been trained as (Place an X adjacent to appropriate response):  X Operator      Trainer

Test Type:	X  Blood Gas     Creatinine   Troponin I

This checklist is designed to assist in establishing training documentation for point-of-care testing per VAMC guidelines.  For complete instructions on the use and care of the i-STAT, please refer to the Operator’s Manual.

ELEMENTS OF COMPETENCY ASSESSMENT KEY:  DO = Direct observation of task performance  DOI = Direct observation of instrument maintenance and function checks  R = Review of QC, patient and proficiency results  A = Performance assessment through previous testing, proficiency results and online re-assessment training and exams  E = Evaluation of problem solving skills M = Monitoring recording and reporting of test results including critical values and actions taken

1  CONTROL TESTING:

Ο   I-STAT:  State VAMC policy for frequency of QC testing: Upon receipt of new lot #, new shipment, or every 30 days. Know when to perform electronic QC for troubleshooting purposes. DO/DOI/E

2  PATIENT TESTING:

Ο Demonstrate knowledge of proper Vacutainer tube selection and sample stability post collection DO

Ο  Demonstrate appropriate sample collection technique, patient identification and specimen labelling DO

Ο  Demonstrate proper application of patient sample to cartridge DO/A

Ο  State VAMC protocol for testing under isolation conditions DO

Ο  Demonstrate proper entry of operator ID by using complete SSN number (i-STAT) DO

Ο  Demonstrate proper entry of patient ID by manual entry of complete SSN or scanning armband DO/A

Ο  Perform a patient test using designated test patient ID.  Simulate using real specimen.DO/R/M

Ο  Demonstrate proper scanning of cartridge, specimen preparation, and cartridge loading.DO/R/M

Ο  Successfully perform patient test or QC test .DO/R/M

Ο  State VAMC ‘action range’ or ‘out of range’ policy and actions to be taken (Provider notification comment with read back) .DO/R/M

3  DATA REVIEW:

Ο  Demonstrate recall patient results from the Data Review option.DO

4  DATA UPLOAD:

Ο  Describe upload/download of i-STAT test results to data management system via docking station.DO/DOI

O  Demonstrate knowledge of where to find lab results in lab tab once posted to CPRS EMR. DO/R/M

5  TROUBLESHOOTING AND MAINTENANCE:

Ο  State action(s) for troubleshooting technical problems (not downloading data, error messages, etc.) .DO/ E

Ο  Demonstrate acceptable cleaning/disinfecting technique . DO

Ο  Identify rechargeable Li battery for i-STAT and look for proper charging and charging complete notifications in meter. .A/ E

____________________________________                _____________________

Employee/Operator	                                                               Date

__Lisa G. Lee MT (ASCP) ATC__________________________________                 _____________________

Trainer /Reviewer                                                                   Date