## Slide 1: i-STAT Total β-hCG 6-Month Competency
Ancillary Testing
McGuire VAMC
Richmond, VA

## Slide 2: Quality Control
External controls are performed every 30 days:
  Levels 1 and 3 on all meters
  Stored in refrigerator
  Check expiration date
  Mix well
Staff are required to rotate performing QC
Controls are now programmed on a schedule
i-STAT screen will alert you when QC is due

## Slide 3: Specimen collection
Green top tube must be full
Under-filled tubes can cause inaccurate results
Sample is only stable for i-STAT testing with in 30 minutes after collection
Label all specimens with patient label

## Slide 4: Cartridges
There are several areas in this facility that perform Bhcg testing on the i-Stat. The storage requirements are different for different locations (based on preferences).
Please note the following:
  ED – stored on the shelf at room temperature, exp. 14 days after removal from refrigerator
  Cath Lab and IR – stored in the refrigerator
  Warm up time before usage – 5 min for a single cartridge
  Under-filled cartridges cause errors; fill to the fill mark on the cartridge

## Slide 5: Interpretation
Negative: <5.0 IU/L hCG QUAL ( - )
Indeterminate: 5.0 – 25.00 IU/L hCG QUAL ( )  consult with provider
Positive: >25.00 IU/L hCG QUAL ( + )

## Slide 6: Limitations
β-hCG results should always be used in conjunction with other data, e.g., patient’s medical history, symptoms, results of other tests, clinical impressions, etc. The i-STAT β-hCG test results should always be used and interpreted only in the context of the overall clinical picture.
Interfering substances (such as heterophilic antibodies, non-specific proteins, or hCG-like substances) may falsely depress or elevate results. In cases where test results are inconsistent with clinical information, results should be confirmed by an alternate hCG method.

## Slide 7: Limitations cont.
Grossly hemolyzed samples can cause a decreased alkaline phosphatase activity, resulting in decreased detection of hCG or quality check codes.
Specimens from peri- or post-menopausal women may elicit weak positive results due to low β-hCG levels unrelated to pregnancy. With a weak positive result, it is good laboratory practice to resample and retest after 48 hours.