## Slide 1: i-STAT Total β-hCG 1st Annual Competency
Ancillary Testing
February 2017
McGuire VAMC

## Slide 2: Quality Control
External controls are performed every 30 days:
  Levels 1 and 3 on all meters
  Stored in refrigerator
  Check expiration date
  Mix well
Staff are required to rotate performing QC
Controls are now programmed on a schedule
i-STAT screen will alert you when QC is due

## Slide 3: Quality Control cont.
Only select Option 2: Schedule 1
Due date: 2nd Tuesday of the month @ 0300
Grace period: 48 hours
Can not be performed before the due date
Analyzers will lock if QC is not performed within the grace period - you will NOT be able to perform patient testing!

## Slide 4: Specimen collection
Green top tube: must be full
Under-filled tubes can cause inaccurate results
Sample is only acceptable for i-STAT testing up to 30 minutes after collection
Label all specimens with patient label

## Slide 5: Cartridges
β-hCG cartridges should only be stored in the med room on the counter.
They should NOT be moved to the i-STAT station!!
This will prevent testing the wrong cartridge.
**Be aware of the cartridge you have picked up!

## Slide 6: Limitations
β-hCG results should always be used in conjunction with other data, e.g., patient’s medical history, symptoms, results of other tests, clinical impressions, etc. The i-STAT β-hCG test results should always be used and interpreted only in the context of the overall clinical picture.
Interfering substances (such as heterophilic antibodies, non-specific proteins, or hCG-like substances) may falsely depress or elevate results. In cases where test results are inconsistent with clinical information, results should be confirmed by an alternate hCG method.

## Slide 7: Limitations cont.
Grossly hemolyzed samples can cause a decreased alkaline phosphatase activity, resulting in decreased detection of hCG or quality check codes.
Specimens from peri- or post-menopausal women may elicit weak positive results due to low β-hCG levels unrelated to pregnancy. With a weak positive result, it is good laboratory practice to resample and retest after 48 hours.

## Slide 8: Troubleshooting
When repeated cartridge failures, failed internal QC checks or suppressed patient results occur and cannot be resolved through repeat testing, notify Ancillary Testing at x5003 or x5885.
If there are problems with analyzer function, use another analyzer and notify Ancillary Testing immediately.

## Slide 9: Interpretation
Negative: <5.0 IU/L hCG QUAL ( - )
Indeterminate: 5.0 – 25.00 IU/L hCG QUAL ( )  consult with provider
Positive: >25.00 IU/L hCG QUAL ( + )