## Slide 1: First annual i-STAT: Troponin Competency
Ancillary Testing
McGuire VAMC

## Slide 2: Ancillary Testing Policies and Procedures
Located on RIC SharePoint
RIC SharePoint icon is located on the main screen of your PC
Use the link below to access the SOP. https://vaww.visn6.portal2.va.gov/apps/richmond/plms/PLMS%20SOPs/POC.53%20%20i-STAT%20Troponin%20I%20Procedure.docx

## Slide 3: Cartridges
Once warmed to room temperature, cartridges are good for 14 days
Handle cartridges by sides only
Don’t press on reagent bubble on top
Close sample well completely before inserting into meter

## Slide 4: cTnI cut-off
Cut-off is now 0.10 ng/mL
Values > 0.10 are considered abnormally elevated for Troponin on i-STAT
Please note that values above the cut-off are considered abnormal but may NOT be consistent with an ischemic event.

## Slide 5: Serial Testing
ONE elevated troponin result does not mean NSTEMI or an acute event.
Serial troponin testing is REQUIRED to determine if patient troponin levels are rising/falling/stable.
Current guidelines support 3-hour serial testing protocols.

## Slide 6: Specimen collection
Green top tube must be full and mixed well
Under-filled tubes can cause inaccurate results
Hemolysis falsely decreases troponin levels
  Good blood collection technique is CRITICAL!
Label all specimens with patient label
Retain ALL troponin samples until ancillary testing picks them up

## Slide 7: cTnI Testing
Sample is only acceptable for i-STAT troponin testing up to 30 minutes after collection
Do not overfill the cartridge
Avoid air bubbles when applying sample
Do not place i-STAT on the docking station while testing is occurring
Do not move the meter once cartridge is loaded

## Slide 8: Quality Control
External controls are performed every 30 days:
  Levels 1 and 3 on all meters
  Stored in refrigerator
  Check expiration date
  Mix well
Staff should rotate performing QC
Controls are programmed on a schedule
i-STAT screen will alert you when QC is due

## Slide 9: Quality Control cont.
Only select Option 2: Schedule 1
Due date:1st Tuesday of the month @ 0300
Grace period: 48 hours
Can not be performed before the due date
Instruments will lock if QC is not performed within the grace period = will NOT be able to perform patient testing!

## Slide 10: Cleaning
You MUST clean the i-STAT after each use with SaniWipes.
Cleaning is a Joint Commission requirement
Do not place medications, food or trash around or near the meters and docking stations.

## Slide 11: Docking/Transmission
When meters successfully transmit data, the arrows on the screen will rotate in circles
If the arrows do not appear or rotate, do the following:
  Check data cable and battery cable
  If cables are connected, dock meter on another station, then….
  Call Ancillary Testing ASAP (x5003, x5885)