## Slide 1: Hematology Competency

## Slide 2: Test Performance
All tubes should be labeled with:
  Full name
  SS#
  Date
  Time drawn
  Phlebotomist/nurse initials
Check that the original label matches the accessioning label prior to testing.

## Slide 3
At least 1mL of EDTA-K2 blood, is needed to automatically aspirate from the Sysmex XN-9000 and the Tosoh G8.
A full 0.3mL tube of 3.2% sodium citrate blood, must be collected to perform testing on the Stago Compact.

## Slide 4
At least 500µL of EDTA-K2 blood is needed to perform testing on the iSED analyzer.
At least 800µL of 3.2% sodium citrate blood is needed for testing on the PFA analyzer.

## Slide 5
Causes of specimen rejection include:
  mislabeled specimens
  insufficient volume of blood
  expired collection tubes
  improper specimen collection
  clotted specimens
Providers should be notified of all rejected specimens.

## Slide 6: Recording/Reporting Results
Results are reviewed for accuracy and comparison with previous results before verification via the delta check policy for hematology.
Critical values are always called to the patients’ doctor or designee in a timely manner.

## Slide 7
The elements required for proper critical value notification include ALL of the following:
- canned comment “CRIT”
- clinicians full name and title
- turn around time
- time of call
- initials of the technologist

## Slide 8: Review of Intermediate Test Results or Worksheets
Quality Control should be performed and documented on the appropriate log sheets.
XN Check is performed once every shift and is stable for 7 days.
XN Check BF is performed once on day shift and is stable for 30 days.

## Slide 9
STA-Coag Control N+ABN is performed once per shift and is stable for 8 hours after reconstitution.
STA-Liatest Control N+P is performed once per shift as needed with the receipt of a D-Dimer test and is stable for 8 hours after reconstitution.

## Slide 10
Tosoh QC material is performed once on day shift and after reconstitution, is stable for 60 days when stored tightly capped at 2°- 8°C.
iSED QC is performed once a day and once opened, is stable for 31 days.
The Self Test procedure for the PFA analyzer is performed once on day shift.

## Slide 11
Sperm-Chex Post VC is performed once on day shift, each day of patient testing and is stable for 45 days once opened.
A peripheral smear containing eosinophils is stained and recorded each day of patient testing as QC for the urine eosinophil samples.
Crystal examination QC is reviewed each day of patient testing.

## Slide 12
Document all (daily, weekly, monthly, and as needed) maintenance procedures on the appropriate log sheets.
MTS/TMS tests that are assigned to the technologist(s) must be completed by the designated time frame.

## Slide 13: Instrument Maintenance
Perform and document daily Sysmex XN and as needed RU-20 maintenance on the XN QC Log.
  When replacing reagents, be sure to add your initials to the comment box onboard the Sysmex.
Perform and document daily, weekly, as needed maintenance, and reagent replacement on the SP-10 Maintenance Log.
Perform and document daily, weekly, monthly, and as needed maintenance on the Stago Compact Maintenance Log.

## Slide 14
Perform and document daily and as needed maintenance on the Tosoh Maintenance Log.
Perform and document daily, weekly, and yearly maintenance on the PFA-100 QC and Maintenance Log.
Perform and document daily and other maintenance on the Hematek Stainer Log.

## Slide 15: Assessment of Test Performance
The hematology department is enrolled in the proficiency-testing program through CAP.
Specimens for proficiency testing are integrated into the routine workload, and are tested by technologists who routinely test patient samples.
Replicate analysis of survey samples is acceptable only if patient samples are routinely analyzed in the same manner.

## Slide 16
All proficiency surveys should be completed within the specified time given in the survey instructions by the assigned technologist(s).
Survey results are reviewed, initialed and dated by the Chief of PLMS, and/or the Director of Hematology and the hematology supervisor.

## Slide 17: Evaluation of Problem Solving Skills
Corrective action of QC results
1. Check the reagent levels and control volume then rerun control.
2. Rerun fresh vial of control and/or reagent.
3. Perform maintenance and cleaning of affected parameters.
4. Contact the supervisor and the appropriate customer technical support center for further instructions.

## Slide 18
Once corrective action has been taken and QC is within the expected range, fill out the action log as follows:
- note specifically which control level and parameter was out
- action taken
- outcome
- date
- initials

## Slide 19
For troubleshooting, see the Troubleshooting section of the Operator’s Manual for the respective analyzer.
If all troubleshooting efforts have been exhausted, place a service call and fill out the action log as follows:
- note the issue
- action taken
- outcome
- reference number (if applicable)
- date
- initials

## Slide 20: You have successfully completed the Hematology Competency for FY18!