## Slide 1: Clinitek status plus ®
Annual competency
Must score 80% or more

## Slide 2
Purpose:
To establish consistent reading of urinalysis test in a point-of-care environment. The test is intended for use for patients experiencing a loss of kidney function, urinary track infections, carbohydrate disorders, or liver function problems.

## Slide 3: Principle:
Clinitek status plus ® is a semi-automated instrument designed to test urine Bilirubin, Blood (Occult), Glucose, Ketones Leukocytes, Nitrites, PH, Protein, Specific Gravity, and Urobilinogen.

## Slide 4: Specimen
Collect freshly-voided urine in a clean container and test as soon as possible. The container should be filled with enough urine to allow for complete dipping of the reagent test strip areas.
  Samples are stable for 1 hour at room temperature and 4 hours refrigerated. NOTE: Sample must be at room temperature prior to analysis.

## Slide 5: Quality control
Must be performed every 24 hour or before patient testing if it is not done daily.
Uses Dropper plus urinalysis dipstick controls 1 and 2
Please use ‘Strip QC” from the display.

## Slide 6: Post analysis
Once the testing is complete, result will print.
Make sure to use full patient ID in order the result to transmit in to CPRS.
No order is needed for urinalysis; and please do not us order number when entering patient ID.

## Slide 7: Quality control troubleshooting
Check for:
  Expiration date for Control solutions
  Expiration date for reagent strip
  If QC fail 2X, repeat with new vials of controls and if fail again, notify Ancillary department

## Slide 8: Limitations:
The following substances in urine may affect the readability of the test pad.
  Visible blood
  Drugs such as Azo
  Nitrofurantoin
  Riboflavin

## Slide 9: Maintenance
On daily basis, clean the test table insert with wet gauze
Clean outside with wipe
Ensure the analyzer operates properly
Note: Wear personal protective equipment. Use universal precautions.

## Slide 10: Weekly cleaning
1. Remove test table by pulling it slowly out of the analyzer
2. Lift the table insert from the test table
3. Drain the drip tray if necessary
4. Wet a cotton tipped stick with water and thoroughly scrub the test table and table insert, Except the white calibration bar
5. Rinse both sides of the table insert and test table under running water.
6. Dry the test table insert, examine the white calibration bar, and insert the test table slowly halfway to the instrument.
If the white calibration bar is dirty, clean with soft cloth or lint free tissue.

## Slide 11
To access RIC SharePoint for the hospital policy and procedure use the following link.
https://vaww.visn6.portal2.va.gov/apps/richmond/plms/PLMS%20SOPs/Forms/Current.aspx
And select “point-of-care” from the list. Use the link “Point of Care SOPs”
Anew website/medialab opens and the “Point Of Care/Ancillary” link will expand. Look for the correct instrument SOP and review.
You have completed this portion . Please go to the exam section. Passing score is 80% or above.