## Slide 1: i-STAT Total β-hCG 1st Annual Competency
Ancillary Testing
McGuire VAMC

## Slide 2: Quality Control
External controls are performed every 30 days:
Levels 1 and 3 on all meters
Stored in refrigerator
Check expiration date
Mix well
Staff are required to rotate performing QC
Controls are now programmed on a schedule
i-STAT screen will alert you when QC is due

## Slide 3: Quality Control continued
Only select Option 2: Scheduled
Due date: 1st Tuesday of the month
Grace period: 48 hours for ED; 96 hours for other clinics.
Can not be performed before the due date
Analyzers will lock if QC is not performed within the grace period - you will NOT be able to perform patient testing!

## Slide 4: Specimen collection
Green top tube: must be full
Under-filled tubes can cause inaccurate results
Sample is only acceptable for i-STAT testing up to 30 minutes after collection
Label all specimens with patient label

## Slide 5: Cartridges
ED: β-hCG cartridges are stored at room temp by the i-stat area
**Other clinics: stored in the Cath-lab refrigerator. Make sure to let it sit at room temp for 5 minutes before use.

## Slide 6: Limitations
β-hCG results should always be used in conjunction with other data, e.g., patient’s medical history, symptoms, results of other tests, clinical impressions, etc. The i-STAT β-hCG test results should always be used and interpreted only in the context of the overall clinical picture.
Interfering substances (such as heterophilic antibodies, non-specific proteins, or hCG-like substances) may falsely depress or elevate results. In cases where test results are inconsistent with clinical information, results should be confirmed by an alternate hCG method.

## Slide 7: Limitations continued
Grossly hemolyzed samples can cause a decreased alkaline phosphatase activity, resulting in decreased detection of hCG or quality check codes.
Specimens from peri- or post-menopausal women may elicit weak positive results due to low β-hCG levels unrelated to pregnancy. With a weak positive result, it is good laboratory practice to resample and retest after 48 hours.

## Slide 8: Troubleshooting
When repeated cartridge failures, failed internal QC checks, or suppressed patient results occur and cannot be resolved through repeat testing, notify Ancillary Testing at x5003, x5885, or x3305.
If there are problems with analyzer function, use another analyzer and notify Ancillary Testing immediately.

## Slide 9: Interpretation
Negative: <5.0 IU/L hCG QUAL ( - )
Indeterminate: 5.0 – 25.00 IU/L hCG QUAL ( )  consult with provider
Positive: >25.00 IU/L hCG QUAL ( + )