## Slide 1: 2021 Annual i-STAT ACT-k Competency

## Slide 2: Items Required
Read this review
Take the online exam and score ≥80%
Contact Ancillary at ext. 3305 or 5885 to schedule a practical – includes testing a liquid control and previously tested sample. Each person must perform the tests individually.

## Slide 3
Locating Ancillary Testing Procedures
Ancillary Testing Policies and Procedures ARE FOUND IN SOFTWARE CALLED “MEDIALAB FEDERAL” Use the link below to access the SOP. \\v06.med.va.gov\ric\service\LaboratoryAdministrativeManual 1. Open the POC Ancillary Backup 2. Open “I-stat” 3. There is document for each test

## Slide 4: i-STAT ACT Assay
Purpose: method of choice for monitoring heparin therapy
Patient Identification: Use a minimum of 2 patient identifiers
    Full name
    Full social security number
    Date of birth
Specimen: whole blood drawn in a plain, un-heparinized plastic syringe.

## Slide 5: i-STAT ACT-k Cartridges
Refrigerated – stable until manufacturer’s expiration date printed on the wrapper.
Must be brought to room temperature at least for 5 minutes before use
Room Temperature – stable for 14 days. Each cartridge must be labeled properly with new expiration date when removed from fridge.
Must close the sample well after sample application and prior to insertion into i-STAT

## Slide 6: Quality Control
  Liquid QC is performed every 30 days
  The i-STATs are programmed with a QC schedule, the screen of the meter will alert you when QC must be performed.
  If QC has not been performed by the date/time on the screen, the meter will not allow patient testing to be performed until QC is completed.

## Slide 7: Performing Quality Controls
  Re-constitute liquid controls according to the manufacturer's directions (one level at a time)
  Run each control level immediately after re-constitution
      Select "Menu"
      #3 - Quality Tests
      #1 – Control
      #2 – Schedule 1
      #1 – i-STAT ACT
      Select appropriate level
      Enter/Scan Operator ID
      Scan Control Lot Number
      Scan Cartridge Lot Number
      Apply the control material to the cartridge
      Immediately close the cartridge and insert it into the analyzer
  Upon completion of test, screen will show Pass/Fail
  If QC fails, repeat the test with freshly re-constituted control solution

## Slide 8: Patient Results
Assay reportable range: 50-1000 sec.
Results showing *** should be repeated with a fresh sample and a new cartridge.
Results >1000 sec should be repeated with a fresh sample and a new cartridge.

## Slide 9
Maintenance
Clean
Clean the i-STAT analyzer with purple top Sani-Wipes after each test. Dwell time is 2 minutes. (Sanitizer liquid must remain on surface for 2 minutes in order to properly sanitize.)
Keep
Keep the i-STAT on the docking station to re-charge the battery, transfer patient results to the electronic record, and for software updates.
If onscreen arrows do not rotate when docked, notify Ancillary

## Slide 10: Sources of Error and Limitations
  Poor sample collection technique
  Delay in testing
  Under/overfilling sample well
  Movement of analyzer during testing
    Error codes may result if the analyzer is moved during testing
    Results may be affected more than ±10%

## Slide 11
What’s Next?
Take the online exam. Check your email for a message from notify@medtraining.org that contains a link to the exam and accompanying PowerPoint. Passing is ≥80%
Contact Ancillary to arrange a time to complete your practical
Please contact Ancillary Testing if you have any questions. Ext. 3305, 5885, or 5003