## Slide 1: 2021 Annual GEM 5000 Blood Gas Competency

## Slide 2
Competency Requirements:
Review this PowerPoint
Take the online exam and score ≥80%
Complete a practical test under observation
  Contact Ancillary at ext. 3305, 5885, or 5003 for availability

## Slide 3: GEM 5000 Blood Gas Review
The GEM 5000 analyzer uses cartridge-based technology to aid in the diagnosis of acid/base status, oxygen delivery capacity, and electrolyte and metabolite balance.
IQM2 (Intelligent Quality Management) provides continuous quality control monitoring of the analytical process before, during, and after sample measurement
Uses 5 process control solutions (PCS) performed continuously each day to confirm sensor, CO-OX, and PAK performance
Uses “IntraSpect” technology to detect abnormal sensor response or absorbance residual error due to micro clots, microbubbles, and/or interfering substances
  If any measurement value is outside the allowable limits, patient sample results for the affected analyte will not display on the sample printout
  If corrective action is not successful, iQM2 will automatically disable the affected analytes(s), making the analyte unavailable for patient analysis

## Slide 4
APV (Auto Pack Validation)
    Used to complete the calibration process to ensure accurate results
    Cartridge warm-up: 45-60 min
    APV must pass before patient samples can be tested.
    APV will be performed once cartridge warm-up is complete. If any analyte fails on the first run, the APV will run again to correct the failed analyte. If APV fails more than twice, the cartridge should be changed.
    Please let Ancillary know if a cartridge needs to be replaced due to errors or APV does not pass multiple times

## Slide 5
Analyzer Test Menu
Measured tests available on the GEM 5000 analyzer:
The following derived (calculated) parameters may be available on an analyzer, depending on the analyzer configuration:
  pH, pCO2, pO2, sodium, potassium, ionized calcium, glucose, lactate, Cl¯, tHb, O2Hb, COHb, MetHb, and HHb
HCO3¯, BE(B), tCO2, HCT, sO2, CaO2, and AG

## Slide 6
GEM Cartridges
Each cartridge contains reagents, controls, and a sensor card where testing is performed
Eliminates operator maintenance for up to 30 days
Cartridges are stored at room temperature (15 - 25°C)
Cartridge Expiration:
  In foil: Cartridge may be inserted up to and including the use-by date printed on the foil
  On-board:
  Cartridge must be replaced when it runs out of tests or when cartridge use-life is reached
  Once the cartridge is removed, it cannot be put back on the analyzer

## Slide 7: Troubleshooting
How to Manually Perform an iQM2 Process:
  Main Menu  Diagnostics  Run iQM2
When to initiate an iQM2 Process:
  If an operator notices frequent messages regarding clots or interferences
  If there are delays with the APV (Auto Pack Validation) process.
What is displayed:
  “No Errors” message is displayed on print out when all issues have been cleared
  The target analyte will be permanently disabled if iQM2 is unable to resolve the issue.

## Slide 8: Incalculable results
Analytes with measured values outside the Reportable Range are reported with a > or < symbol. “Incalculable” will be displayed for results that are outside the measuring capability of the analyzer. Often, this occurs with interferences.
When this happens:
Discard results
Re-collect fresh sample
Re-test

## Slide 9
ABG Specimens
Collection Device: Use ONLY lyophilized lithium heparin blood gas syringes with a concentration of ~ 25 USP units/mL
Specimen Types
  Arterial:
    Steady state of ventilation for 30 minutes
    Use the radial, brachial, or femoral artery
    Document location of sample draw and the presence of collateral blood flow (Allen’s test) on the instrument itself
  Venous:
Antecubital vein preferred
Not suitable for oxygenation studies

## Slide 10
Specimen Handling Tips for Reliable Results
ONLY use ABG syringes
Always discard 3-5 cc from a line before specimen is collected
Expel air after collection and cap syringe
Mix immediately at the bedside: rock x5 and roll x5
Label syringe at bedside
Mix again at the analyzer: rock x5 and roll x5
Sample must be tested within 30 minutes of collection. However, Lactate will not be accurate if tested after 5 minutes.**
Expel air
Rock and roll

## Slide 11
  Effects of Poor Sample Handling
  Micro-clots form when sample is not mixed immediately after collection and again at the analyzer
  pO2 ↑ when air is not expelled prior to mixing
  Lactate ↑ when sample analysis is delayed more than 5 minutes after collection
  HCT ↑or↓ when adequate mixing is not performed prior to running a specimen
  Electrolytes/Glucose results may vary with test delays >30 minutes
  Do not place blood gas specimens on ice
  Extended application of tourniquet, clinching fist, drawing too forcefully with the syringe, and using a smaller needle during sample collection could cause hemolysis and K+ will be falsely elevated

## Slide 12
Never test body fluids on any of the IL GEM blood gas analyzers through point- of-care testing!
If a provider requests a pH on a body fluid, please refer them to the Main Lab Chemistry Dept. (3A-138) at x5722, option #3

## Slide 13
Processing Samples
  Select sample type from the Quick Start screen
  Enter or scan Operator ID
  Mix sample and insert probe into syringe
  Select “Start Aspirating”
  Remove sample when prompted by audible beep and blinking light
  Enter or scan barcoded Patient ID
  Input the temperature, vent settings, and Allen’s test, as needed on the specimen information screen
  Press “View Results,” or results will show in ~45 seconds
  Select “ACCEPT” to have results transfer to the patient’s record and receive a printout
  Press “DISCARD” if results are unacceptable, invalid, or should not be reported
  Always ACCEPT or DISCARD. Do not leave results PENDING.
  Return to the Quick Start screen

## Slide 14
Patient Results
Results include measured, derived, temp-corrected, and CO-Oximetry
Color-coded test results inform you at a quick glance the status of your patient
  Green lettering on white background is normal
  Yellow background is outside the reference range, but not critical
  Red background is critical
Patient history can be obtained from the results screen at the time the sample is tested, by selecting the “Patient History” button
The reference range will be printed with the patient results

## Slide 15
Trouble-shooting
Perform 2-3 manual I QM processes (Menu  Diagnostics Perform iQM2 Process)
  When the analyzer is posting iQM2 errors or “fixing” frequently
  When you receive multiple analyte failures (APV failures) just after a new cartridge has been placed on the analyzer.
If iQM2 processes do not resolve, replace the cartridge
Use another blood gas analyzer from another unit
Place an order and send a labled sample to the Main Lab
Notify Ancillary Testing when you can’t resolve the issue
    Isabel x5003
    Mulu x5885
    Heather x3305

## Slide 16
Ancillary Testing Policies and Procedures ARE FOUND IN SOFTWARE CALLED “MEDIALAB FEDERAL” Use the link below to access the SOP. \\v06.med.va.gov\ric\service\LaboratoryAdministrativeManual 1. Open the POC Ancillary Backup 2. Open Point of Care Ancillary 3. Open Blood Gast Testing

## Slide 17
What’s Next
Complete the online quiz
Contact Ancillary to complete your practical at Ext. 3305, 5885, or 5003
Or see one of the nurse educators at your unit.
Thank you!