## Slide 1: Uscreen UDS Cup 1st Annual Competency

## Slide 2: Competency Requirements
Review this PowerPoint
Complete the online exam
Passing is an 80% or higher
Perform controls

## Slide 3: Overview
UScreen® Drugs of Abuse Cup is a competitive immunoassay that is used to screen for presence of various drugs in urine. It is a chromatographic absorbent device in which drugs in a urine sample competitively combine to a limited number of drug monoclonal antibody conjugate binding sites. When the test is activated, the urine is absorbed into each test strip by capillary action, mixes with the respective drug monoclonal antibody conjugate, and flows across a pre-coated membrane. When the drug levels within the urine sample is below the detection level of the test, respective drug monoclonal antibody conjugate binds to the respective drug-protein conjugate immobilized in the Test Region (T) of the test strip. This produces a colored test line in the Test Region (T) of the strip that regardless of its intensity indicates a negative result.
When the sample drug levels are at or above the detection level of the test, the free drug in the sample binds to the respective drug monoclonal antibody conjugate, preventing the respected drug monoclonal antibody conjugate from binding to the respective drug-protein conjugate immobilized in the Test Region (T) of the device. This prevents the development of a distinct colored band in the test region, indicating a preliminary positive result

## Slide 4: You will need
Uscreen UDS Cup
Fresh urine specimen
2 levels of controls
Timer
PPE

## Slide 5: reagents
Uscreen UDS cups are stored and used at room temperature. UDS external controls are refrigerated.
Cups are good until the expiration date printed on packaging

## Slide 6: Quality controls

## Slide 7: Internal procedural Controls
Verifies the assay validity (performance of the cup)
A line at the “control” position indicates the sample migrated properly and that the reagents worked
If the line at the “control” position is not present, the test is invalid. DO NOT report patient results. Repeat the test with a new sample and new cup.

## Slide 8: External Positive & negative Controls
Verifies the performance of the reagents and the operator.
External controls must be performed monthly.
External controls are performed by the same process as patient testing,
If the correct control results are not obtained, do not report patient results. Repeat the test with a new control samples and a new cup. If the correct results are still not obtained, do not continue patient testing, contact Ancillary at ext. 3305, 5885, or 5003.

## Slide 9: Specimen collection
Remove the cup from the sealed pouch.
Label the cup with patient’s full name, social security, date and time of collection.
Instruct the patient to collect an adequate sample volume of urine (30 mL minimum) in the cup and secure the cap tightly.
**Treat all specimens as potentially infectious. Follow universal precautions**

## Slide 10: Patient testing
Read the urine temperature immediately (as soon as possible) after the patient has voided. Temperature must be between 90-100F.
Place the cup on a flat surface and peel off the label on the cup to view the results.
Read the adulterant – creatinine test result within 5 min; see the color chart for interpretation. If the specimen indicates adulteration, do not interpret the drug results and collect a new sample.
Read the test results at 5 minutes. Results read after more than 5 min may not be accurate and should not be read.
The intensity of the control and test line may vary. Any line, without regard to intensity or size, should be noted.
Document results in WRX

## Slide 11: Interpreting results
Preliminary Positive
No line appears in the test zone (T). If no color band appears in the appropriate test “T” region, a preliminary positive result is indicated for the corresponding drug of that specific test zone.
Preliminary Negative
One colored line appears in the Control zone and another line appears in the Test zone (T). The negative result indicates that the concentration of the corresponding drug of that specific test zone is absent or below the detection limit of the test.
Invalid
If no C line or no C line and T line develop within 5 minutes, the test is invalid. Repeat the test

## Slide 12: limitations
DO NOT USE this device to test substances other than urine
Adulterated urine samples may produce erroneous results. Strong oxidizing agents such as bleach can oxidize drug analytes.
This test is a qualitative screening assay. It is not designed to determine the quantitative concentration of drugs or the level of intoxication

## Slide 13: The UScreen UDS SOP can be found at the following link: \\v06.med.va.gov\ric\service\LaboratoryAdministrativeManual
Select “POC Ancillary Backup”
Select “Point of Care Ancillary”
Select “Manual Tests”
Locate the Uscreen UDS Procedure

## Slide 14: What’s next?
Take the online exam – passing is an 80% or higher
1
Complete controls and enter in WRX
2
Thank you!
3