## Slide 1: 2023 i-Stat Creatinine Annual Competency
Competency includes:
Presentation
MTS test, minimum 80% for passing
Practical Portion- Test the external liquid controls

## Slide 2: Purpose
Elevated levels of Creatinine are associated with abnormal renal function and will occur when there is a significant reduction of the glomerular filtration rate or when urine elimination is obstructed. Creatinine is a better indicator of renal function than urea because it is less dependent on diet, exercise and hormones.

## Slide 3: QUALITY CONTROL (QC)
2 levels; performed every 30 days
i-STAT screen will alert when due
Scheduled for 1st Tuesday of each month
Meter will lock out patient testing if QC is not performed by the end of the grace period (72 hours)
Operators must rotate; see schedule from Ancillary Testing via Outlook e-mail

## Slide 4: Specimen
A specimen can be collected in one of two different ways:
  Venipuncture: lithium or sodium heparin collection tubes.
  Plain syringes are acceptable, but sample must be TESTED IMMEDIATELY.
  Arterial puncture: blood gas syringe with heparin
    Syringes or tubes are obtained from SPD and stored at room temperature.

## Slide 5: REAGENT
i-STAT Creatinine Cartridges
    Refrigerated cartridges at a range of 2-8°C are stable until the expiration date indicated on the box
    Room Temperature cartridges can be stored at 18 - 30°C and are stable for 14 days. Individual cartridges must be re-dated to indicate the 14-day expiration.
    Cartridges should warm to room temp before performing testing
    **ONCE AT ROOM TEMPERATURE, THE CARTRIDGES CAN NOT BE RETURNED TO THE REFRIGERATOR; THEY MUST BE RE-DATED WITH A 14-DAY EXPIRATION.

## Slide 6: Procedure
  Turn the i-STAT on
  Select option #2: i-STAT cartridge
  Enter your operator ID
  Identify patient ID (use minimum two patient identifiers)
  Scan lot number of cartridge
  You have 15 minutes to collect the sample
  Take the cartridge out of the pouch
  Dispense a drop of blood into the cartridge making sure it fills up to the fill mark
  Close the cover over the sample well
  Insert the cartridge into the port on the analyzer; note: never attempt to remove a cartridge while the “Cartridge Locked” is displayed

## Slide 7: Troubleshooting
Calibration is automatically performed as part of the test cycle on each cartridge
Repeat any questionable result or if there is an instrument error. You may see the following: “***”, failed internal QC Check. Ancillary Testing should be notified at x5885 or x3305 when failed internal QC or repeated “**” error codes observed.
Note: If testing a sample collected in a plain syringe, when repeating you must use a freshly collected sample.

## Slide 8: Interpretation
Results are displayed with their units and depicted as bar graphs with reference ranges marked under graphs.
EXPECTED VALUE: 0.6 - 1.3 mg/dL
REPORTABLE RANGE: 0.2 – 20.0 mg/dL
For administration of Contrast Media based on eGFR values, see your department’s most recent “Contrast Media Guide”. The laboratory implemented a new eGFR formula on March 1st, 2022.

## Slide 9: Limitations/Interfering Factors
An interferent is a substance which, if present at significant levels in the blood specimen being analyzed, will produce an error in the result of the analyte being measured. The following table identifies known interfering substances based on the manufacturer data on i-STAT creatinine cartridge.

## Slide 10: Limitations/Interfering Factors
Creatinine | Interferent | Interferent Conc. | Effect on Creatinine
 | Acetaminophen | 1.32 mmol/L | Toxic level; increase (↑) creatinine
 | Acetylcysteine | 10.2 mmol/L | Used to reverse acetaminophen toxicity; increase (↑) creatinine
 | Ascorbate | 0.34 mmol/L | Increase (↑) creatinine by 0.3 mg/dL
 | Bromide (therapeutic) | 2.5 mmol/L | Associated with halothane anesthesia; Increase (↑) creatinine
< 2 mg/dl | pCO2 | Above 40 mmHg | Increase (↑) creatinine by 6.9% per 10 mmHg PCO2
 |  | Below 40 mmHg | Decrease (↓) creatinine by 6.9% per 10 mmHg PCO2
>2 mg/dl | pCO2 | Above 40 mmHg | Decrease (↓) creatinine by 3.7% per 10 mmHg PCO2
 |  | Below 40 mmHg | Increase (↑) creatinine by 3.7% per 10 mmHg PCO2
 | Glycolic Acid | 10.0 mmol/L | Decreased (↓) creatinine; use another method
 | Hydroxyurea | 0.92 mmol/L | Used in treatment of cancer, sickle cell and HIV infection, polycythemia, thrombocythemia and psoriasis; increase (↑). Use another method
 | Creatine | 0.382 mmol/L creatine | Found in supplements, muscle trauma or myopathies, statins, hyperthyroidism; increase (↑) creatinine by 0.20 mg/d;

## Slide 11: Maintenance/Safety
Point of care devices are considered Reusable Medical Equipment. Make sure the device is cleaned between every patient test. Use the Super Sani-Cloth wipes (purple top) provided by the hospital.
Always use Universal Precautions when performing any testing/procedure using any body fluids.

## Slide 12
You must complete the on-line exam, score >80% and test the external liquid controls.
You have completed the presentation.
Point of Care SOPs are located in Medialab federal. You can access the SOP using these steps:
On the intranet click CVHCS SharePoint Home
Click on Standard Operating Procedures under Exec.Office (first column)
Click on Pathology & Laboratory SOP Site (right side towards bottom)
Click on the link displayed under “For all Point of Care SOPs follow this link to the POC page”, (right side towards bottom)
Contact Ancillary department or your supervisor if you can’t access the POC SOPs.