## Slide 1: Annual Avoximeter Competency
Cardiac Cath Lab
Read this module.
Complete the test following the module; you must score ≥80% to be recertified.
Schedule your Practical with Ancillary.

## Slide 2: Principle
The Avoximeter illuminates the sample with multiple wavelengths of light, records the optical absorbance of the sample at each wavelength, and computes the results.
In 10 seconds, the total hemoglobin concentration and the relative concentrations of oxy-, carboxy-, and methemoglobin are shown in appropriate units on the liquid-crystal display on the front panel.

## Slide 3: Specimen
Heparinized whole blood obtained in a plain syringe by the physician or physician assistant from catheters used in the catheterization lab.
The minimum acceptable sample size is 50 µL.
Sample must be rolled between palms to keep the red cells and plasma well mixed.

## Slide 4: Quality Control Procedure
Quality Control for the Avoximeter is composed of 3 processes:
Setting the path length of test cuvettes
Optical Filters
Aqueous Control Materials
Each package of cuvettes is labeled with a lot Number and Calibration code. Visually verify these match every time a sample is analyzed!

## Slide 5: Optical QC
Each Avoximeter is equipped with analyzer-matched cuvette-shaped yellow and orange optical QC filters that can be inserted into the instrument to simulate a blood sample of known composition.
The manufacturer recommends that quality control testing with both filters be performed:
Once each day of testing
When an analyzer is brought from a hot to a cold temperature, or vice versa
Whenever the performance of the analyzer requires verification

## Slide 6: Liquid QC (Aqueous QC)
The purpose of liquid controls is to verify that the Avoximeter system is functioning properly throughout the testing process.
The manufacturer recommends weekly quality control testing with liquid controls to:
Verify proper storage of cuvettes
Verify proper shipping conditions before a new lot of cuvettes is put into use

## Slide 7: Patient Identification
Before the collection of any type of blood specimen, the patient must be properly identified.
Use at least 2 of these patient identifiers:
Full name
Full social security number
Date of birth

## Slide 8: Patient Testing
Confirm
Confirm the temperature probe is plugged firmly into the rear panel of the instrument.
Verify
Verify that the lot # and calibration code are correct for the cuvettes in use.
Collect
Collect the sample, expel one drop and connect the sample filled syringe to a new disposable cuvette.
Inject
Hold cuvette at 45° angle and inject sample into cuvette until sample reaches the vent patch. Do not overfill – if vent patch is protruding, discard and use a new cuvette. Leave the syringe attached to the cuvette.
Confirm
Confirm that the light path at the widest portion of the sample cuvette is free of debris or air bubbles.
Rotate
Hold the cuvette by the black cap, rotate it so the textured, black side of the vent patch will be on the left when the cuvette is fully inserted into the slot. Must be inserted within seconds of filling it!
Analyze
Results will be displayed after the sample information is entered.
Remove
Remove the cuvette as soon as testing is complete. Press the Enter/On key 2x to return to the Ready screen.

## Slide 9: Ranges
Measurable:
Total Hemoglobin 4-25 g/dL
Oxyhemoglobin 0-100%
Reference (Normal):
Total Hemoglobin
Men 14-18 g/dL
Women 12-16 g/dL
Oxyhemoglobin 92-100%

## Slide 10: Sources of Error/Interfering Factors
Citrate is known to change the pH of the blood and cause errors in spectrophotometric measurements. Similarly, fluoride, oxalate, and EDTA should be avoided.
Excessive volumes of anticoagulant or saline may oxygenate the sample or cause dilution errors.
Cuvettes should be stored in a sealed bag with a desiccant pouch. Desiccant should be checked prior to using the cuvettes to ensure that the color indicator has not changed from blue to pink. If the indicator has changed, do not use the cuvettes until a new desiccant pouch can be added to the bag.
When filling cuvettes, do not insert it into the analyzer if the vent patch protrudes.
If the sample is poorly mixed or contains bubbles, discard the cuvette.

## Slide 11: Maintenance and Safety
Analyzer Cleaning:
Clean the analyzer surface and screen with a soft cloth damped with isopropyl alcohol, distilled water, or Sani Wipes(Purple top)
A cotton swab may be used to clean the outside of the cuvette slot.
Do Not insert the cotton swab into the analyzer.
Safety:
Use universal precautions when performing testing.
All contaminated items should be disposed of in a red/orange lined trash container.
All sharps should be disposed of in sharps containers.

## Slide 12: Questions/Problems
Contact ancillary at 5885, 3305, or 8036
Policy for this test is in Media lab federal. If you need to review this policy and procedure, Use the link below.
\\v06.med.va.gov\ric\service\LaboratoryAdministrativeManual
Click on “POC Ancillary Backup”
“Point of Care Ancillary
“Manual tests”
And select “Avoximeter 4000” Procedure

## Slide 13: What’s Next?
Take the online exam. Check your email for a message from “notify@medtraining.org” that contains a link to the exam and accompanying PowerPoint. Passing is ≥80%
Contact Ancillary to arrange a time to complete your practical.
Please contact Ancillary Testing if you have any questions. Ext. 3305, 5885, or 8036.