## Slide 1: 2023 Annual GEM 5000 Blood Gas Competency

## Slide 2
Competency Requirements:
Review this PowerPoint
Take the online exam and score ≥80%
Complete a practical test under observation
  Contact Ancillary at ext. 3305, 5885, or 8036 for availability

## Slide 3: GEM 5000 Blood Gas Review
The GEM 5000 analyzer uses cartridge-based technology to aid in the diagnosis of acid/base status, oxygen delivery capacity, and electrolyte and metabolite balance.
IQM2 (Intelligent Quality Management) provides continuous quality control monitoring of the analytical process before, during, and after sample measurement
Uses 5 process control solutions (PCS) performed continuously each day to confirm sensor, CO-OX, and PAK performance
Uses “IntraSpect” technology to detect abnormal sensor response or absorbance residual error due to micro clots, microbubbles, and/or interfering substances
  If any measurement value is outside the allowable limits, patient sample results for the affected analyte will not display on the sample printout
  If corrective action is not successful, iQM2 will automatically disable the affected analytes(s), making the analyte unavailable for patient analysis

## Slide 4
GEM Cartridges
Each cartridge contains reagents, controls, and a sensor card where testing is performed
Eliminates operator maintenance for up to 30 days
Cartridges are stored at room temperature (15 - 25°C)
Cartridge Expiration:
  In foil: Cartridge may be inserted up to and including the use-by date printed on the foil
  On-board:
  Cartridge must be replaced when it runs out of tests or when cartridge use-life is reached
  Once the cartridge is removed, it cannot be put back on the analyzer

## Slide 5
APV (Auto Pack Validation)
    Used to complete the calibration process to ensure accurate results
    Cartridge warm-up: 45-60 min
    APV must pass before patient samples can be tested.
    APV will be performed once cartridge warm-up is complete. If any analyte fails on the first run, the APV will run again to correct the failed analyte. If APV fails more than twice, the cartridge should be changed.
    Please let Ancillary know if a cartridge needs to be replaced due to errors or APV does not pass multiple times

## Slide 6
Analyzer Test Menu
Measured tests available on the GEM 5000 analyzer:
The following derived (calculated) parameters may be available on an analyzer, depending on the analyzer configuration:
  pH, pCO2, pO2, sodium, potassium, ionized calcium, glucose, lactate, Cl¯, tHb, O2Hb, COHb, MetHb, and HHb
HCO3¯, BE(B), tCO2, HCT, sO2, CaO2, and AG

## Slide 7: Incalculable results
Analytes with measured values outside the Reportable Range are reported with a > or < symbol. “Incalculable” will be displayed for results that are outside the measuring capability of the analyzer. Often, this occurs with interferences.
When this happens:
Discard results
Re-collect fresh sample
Re-test

## Slide 8
ABG Specimens
Collection Device: Use ONLY lyophilized lithium heparin blood gas syringes with a concentration of ~ 25 USP units/mL
Specimen Types
  Arterial:
    Steady state of ventilation for 30 minutes
    Use the radial, brachial, or femoral artery
    Document location of sample draw and the presence of collateral blood flow (Allen’s test) on the instrument itself
  Venous:
Antecubital vein preferred
Not suitable for oxygenation studies

## Slide 9
Specimen Handling Tips for Reliable Results
ONLY use ABG syringes
Always discard 3-5 cc from a line before specimen is collected
Expel air after collection and cap syringe
Mix immediately at the bedside: rock x5 and roll x5
Label syringe at bedside
Mix again at the analyzer: rock x5 and roll x5
Sample must be tested within 30 minutes of collection. However, Lactate will not be accurate if tested after 5 minutes.**
Expel air
Rock and roll

## Slide 10
  Effects of Poor Sample Handling
  Micro-clots form when sample is not mixed immediately after collection and again at the analyzer
  pO2 ↑ when air is not expelled prior to mixing
  Lactate ↑ when sample analysis is delayed more than 5 minutes after collection
  HCT ↑or↓ when adequate mixing is not performed prior to running a specimen
  Electrolytes/Glucose results may vary with test delays >30 minutes
  Do not place blood gas specimens on ice
  Extended application of tourniquet, clinching fist, drawing too forcefully with the syringe, and using a smaller needle during sample collection could cause hemolysis and K+ will be falsely elevated

## Slide 11
Never test body fluids on any of the IL GEM blood gas analyzers through point- of-care testing!
If a provider requests a pH on a body fluid, please refer them to the Main Lab Chemistry Dept. (3A-138) at x5722, option #3

## Slide 12
Processing Samples
  Select sample type from the Quick Start screen
  Enter or scan Operator ID
  Mix sample and insert probe into syringe
  Select “Start Aspirating”
  Remove sample when prompted by audible beep and blinking light
  Enter or scan barcoded Patient ID
  Input the temperature, vent settings, and Allen’s test, as needed on the specimen information screen
  Press “View Results,” or results will show in ~45 seconds
  Select “ACCEPT” to have results transfer to the patient’s record and receive a printout
  Press “DISCARD” if results are unacceptable, invalid, or should not be reported
  Always ACCEPT or DISCARD. Do not leave results PENDING.
  Verify Patient ID on pop-up screen. If ID not verified, results will not go to CPRS
  Return to the Quick Start screen

## Slide 13
Patient Results
Results include measured, derived, temp-corrected, and CO-Oximetry
Color-coded test results inform you at a quick glance the status of your patient
  Green lettering on white background is normal
  Yellow background is outside the reference range, but not critical
  Red background is critical
Patient history can be obtained from the results screen at the time the sample is tested, by selecting the “Patient History” button
The reference range will be printed with the patient results

## Slide 14
Critical Results
All critical values received on the GEM 5000 must be documented in CPRS using a Critical Value Note
Critical values are indicated:
  The result is highlighted in red on the analyzer screen.
  The result has either ↑ ↑ or ↓↓ next to it on the print-out.
Critical
High/Low
Normal

## Slide 15
Trouble-shooting
Perform 2-3 manual iQM processes (Menu  Diagnostics Perform iQM2 Process)
  When the analyzer is posting iQM2 errors or “fixing” frequently
  When you receive multiple analyte failures (APV failures) just after a new cartridge has been placed on the analyzer.
If iQM2 processes do not resolve, replace the cartridge
Use another blood gas analyzer from another unit
Place an order and send a labelled sample to the Main Lab
Notify Ancillary Testing when you can’t resolve the issue
    Isabel x8036
    Mulu x5885
    Heather x3305

## Slide 16
Ancillary Testing Policies and Procedures ARE FOUND IN SOFTWARE CALLED “MEDIALAB FEDERAL” Use the link below to access the SOP. \\v06.med.va.gov\ric\service\LaboratoryAdministrativeManual 1. Open the POC Ancillary Backup 2. Open Point of Care Ancillary 3. Open Blood Gas Testing

## Slide 17
What’s Next
Complete the online quiz
Contact Ancillary to complete your practical at Ext. 3305, 5885, or 8036
Or see one of the nurse educators at your unit.
Thank you!