## Slide 1: Annual Urine HCG Test Competency

## Slide 2: OSOM® hCG Combo Test
Purpose: It is a rapid immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in urine as an aid in early determination of pregnancy.

## Slide 3: Principle
Urine sample is added to the sample well (S) of the test device. The sample migrates through reaction pads where hCG, if present in the sample, binds to a monoclonal anti-hCG dye conjugate. The sample then migrates across the membrane toward the results window, where the labeled hCG complex is captured at a test line region (T) containing immobilized rabbit anti-hCG . Excess conjugate will flow past the test line region and be captured at a control line region containing an immobilized antibody directed against the anti-hCG conjugate (with or without hCG complexed to it.)

## Slide 4: Specimens and Reagents
REAGENTS:
OSOM® hCG Test Devices individually pouched, each containing a disposable pipette. Stored at room temperature.
UA Controls Abnormal and Normal, stored in a refrigerator (2-8° C)
Clock or Timer
SPECIMEN:
Urine specimen - collected in a clean, dry, plastic container.
May be stored at room temperature at 15°-30°C for up to 8 hours or refrigerated at 2°-8°C for up to 72 hours

## Slide 5: Quality Control (QC)
Internal quality control (procedural control)
A procedural control is incorporated into each OSOM Pregnancy Test Card for routine quality checks. It is required that the control is documented for each sample as part of the sample testing.
The appearance of the control band in the result window is an internal positive procedural control which validates the following:
Test system: assures the detection of the test line and control line is intact, that adequate sample volume was added, and that adequate capillary migration of the sample has occurred. Verifies proper assembly of device.
Operator: ensures the operator added adequate volume of fluid to the sample well. If the control band does not appear at the read time, the test is invalid and the test should be repeated using a new test device.

## Slide 6: Quality Control (QC)
External QC Procedure
  1. Remove the controls from the refrigerator and allow them to come to room temperature.
  2. Remove 2 test devices from the pouches.
  3. Mix the Normal (Green top vial) control gently by inversion to assure homogeneity of the solution.
  4. Remove the control cap and gently squeeze the sides of the bottle. Dispense 4 drops into the sample well labeled with the letter “S” for sample.
  Repeat steps 3 and 4 for Abnormal (Red Top Vial)
  Read the results at 3 mins and record results in WRX.
  **Do not use the bottles after the expiration date!

## Slide 7: Patient Test Procedure
Remove the Test Device and the pipette from the pouch. Place device on a flat surface.
1
Squeeze the bulb of the pipette and insert the barrel into the sample. Release and draw up enough sample to fill the barrel to the fill line.
2
Expel the entire contents of barrel into the sample well of the test device. No drop counting is required. Do not overfill the sample well.
3
Read results at 3 minutes.
4
**Note: Results are invalid after the stated read time.
5

## Slide 8: Interpretation of results
Positive – two separate and distinct bands, at “T” and “C” position
Negative – if no band appears at “T” position and a distinct band is visible at “C” position
Invalid – if no band appears at the “C” position. The test should be repeated using another device
**Note: the intensity of the “test” band may vary with different specimens

## Slide 9: Expected Values
In normal pregnancy, 20 mlU/mL hCG is reported to be present in urine 2 to 3 days before the first missed menstrual period. The levels of hCG continue to increase up to 200,000 mlU/mL at the end of the first trimester.

## Slide 10: Result Reporting
Urine pregnancy result must be reported in the web-based result entry form WRX (Web Result Xpress)
Link to WRX and access code has been provided during training.

## Slide 11: Limitations
This test should only be used for the qualitative detection of hCG in urine for the early determination of pregnancy
For diagnostic purposes, hCG test results should always be used in conjunction with other methods and in the context of the patient’s clinical information
If the hCG level is inconsistent with, or unsupported by, clinical evidence, results should also be confirmed by using a quantitative serum hCG prior to the performance of any critical medical procedure.

## Slide 12
Universal precautions must be used when collecting and testing body fluid specimens, as well as Quality Control material.
Always refer to the OSOM® hCG Combo Test Procedure in POC Policies and Procedures found on Medialab
 Point of Care/Ancillary  Point of Care SOPs. Located in RIC SharePoint, accessing the SOPs through CVHCS Policy Site. You can access it using the link below.
POC SOPs - All Documents (sharepoint.com)