## Slide 1: Annual i-Stat PT/INR Competency
Competency Consists of:
      PowerPoint Presentation
      MTS Online Test
      Practical- Under observation

## Slide 2: Principle
The i-Stat PT/INR test is a whole blood determination of the prothrombin time and INR used for monitoring oral anticoagulant therapy. The test determines the time required for complete activation of the extrinsic pathway of the coagulation cascade when initiated with thromboplastin.

## Slide 3: Specimen, Reagents, and Supplies
SPECIMEN - Whole blood from a finger stick
Use the first drop of blood from the finger stick (do not wipe it off) – apply directly to the cartridge
REAGENTS/SUPPLIES
i-STAT Analyzer
PT/INR cartridges
lancet device
NOTE: cartridges stored at 2-8°C expire on the manufacturer’s expiration date; cartridges may be stored at 18-30°C for 14 days after removal from the refrigerator.
Once removed from refrigerator, cartridges must be dated with new 14-day expiration date!

## Slide 4: Procedure
  Turn the i-STAT on
  Select i-STAT cartridge
  Scan or enter your operator number
  Identify patient (minimum of two identifiers) and scan /enter the patient’s full SSN ID into the analyzer
  Scan lot number of cartridge
  Note: When the screen displays “insert cartridge” you have 15 minutes to collect the sample.
  Take the cartridge out of the pouch
  Clean finger and prick the side of the finger
  Dispense the first drop of blood into the cartridge making sure it fills up to the fill mark
  Close the cover over the sample well
  Insert the cartridge into the port on the analyzer
  ***NOTE: never attempt to remove a cartridge while the “Cartridge Locked” is displayed.

## Slide 5: Quality Control
External Liquid Controls performed every 30 days in each location by the operators (level 1 and 2 on each i-Stat)
i-STAT will indicate on the screen when QC is due.
If QC is not performed by end of grace period, meter will lock out and patient testing is not available.
Reconstitute liquid controls according to the manufacturer's directions. Controls are test specific. Reconstitute only 1 level at a time.
Test each control level immediately after reconstitution.
If QC fails, repeat the test with freshly reconstituted QC
  Contact Ancillary Testing for subsequent failures at
x8036, x5885 or x3305.
  NOTE: CALIBRATION OCCURS AUTOMATICALLY AS PART OF THE TEST CYCLE ON EACH CARTRIDGE

## Slide 6: LIMITATIONS
Cubicin/daptomycin at clinically relevant levels can cause false prolongation of prothrombin time and elevation of INR results.
PT/INR test is insensitive to fibrinogen concentration between 70 and 541 mg/dL.
PT/INR is insensitive to heparin up to 1.0 U/mL.
Poor technique in sample collection can compromise results, ex. air bubbles.
Error codes may result if the analyzer is moved while testing is in progress.

## Slide 7: Recommendations & ranges
FOLLOW-UP RECOMMENDATIONS:
INR results of >4.0 requires repeat.
INR ≥ 4.5 must be repeated and a sample should be sent to the lab for confirmation of results.
Repeat test when results are inconsistent with patient clinical status or when analyzer fails to give a result.
ASSAY/REPORTABLE RANGE:
INR 0.9-6.0 ***
***Performance characteristics have not been established at INRs above 6.0.
REFERENCE RANGE:
Dependent upon the clinical circumstances of individual patients

## Slide 8: Critical values
CRITICAL RESULT: INR ≥4.5
The procedure for handling a critical value is as follows:
Repeat the test immediately
Send a PT/INR sample to the lab for confirmation
Notify the patient’s provider (as appropriate) using read back / verify
Document in CPRS and include notification time to provider

## Slide 9: Miscellaneous
MAINTENANCE:
Clean outside of analyzer after every patient use with a purple Super Sani-Wipe.
TROUBLESHOOTING:
If there are repeated cartridge failures, failed internal QC checks, or suppressed patient results, notify the Ancillary Testing Department.
If i-Stat does not turn on, check batteries and power supply.
SAFETY:
Universal precautions as described by CDC should be used in performing the testing described in this procedure. All contaminated items should be disposed of in red or orange lined trash containers. All lancets must be disposed of in a sharp's container. For more safety information, refer to the Laboratory Safety manual in the back of Ancillary Testing manual and to your Service-specific safety procedures.

## Slide 10
Point of Care/Ancillary SOPs are in Richmond SharePoint and can be accessed the SOP using the link below.
POC SOPs - Pathology & Laboratory - Home (sharepoint.com)
 You can also contact Ancillary department or your supervisor if you can’t access it.

## Slide 11
Please complete the exam and practical before May 31st
Thank you!
Questions or concerns: contact Ancillary at ext. 3305, 5885, or 8036