.

Technical SOP

| Title | Erythrocyte Sedimentation Rate |  |
| --- | --- | --- |
| Prepared by | Wendell McMillan | Date: 07/06/2011 |
| Owner | Altreisha Foster | Date: 07/06/2011 |
| Laboratory Approval | Local Effective Date: |  |
| Print Name and Title | Signature | Date |
|  | Refer to the electronic signature page for approval and approval dates. |  |
| Annual Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	4

5.	Calibrators/Standards	4

6.	Quality Control	4

7.	Equipment And Supplies	5

8.	Procedure	5

9.	Calculations	8

10.	Reporting Results And Repeat Criteria	8

11.	Expected Values	9

12.	Clinical Significance	9

13.	Procedure Notes	9

14.	Limitations Of Method	9

15.	Safety	10

16.	Related Documents	10

17.	References	10

18.	Revision History	10

19.	Addenda	11

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| Erythrocyte Sedimentation Rate | ESR STAT PLUS | ESR |

Synonyms/Abbreviations

Sed Rate, ESR

Department

Hematology

# ANALYTICAL PRINCIPLE

The ESR STAT PLUS analyzer uses centrifugation and laser optic principles to measure the ESR in anticoagulated whole blood.  The ESR STAT PLUS obtains the result in approximately 4 minutes and uses a small volume of blood.  A filled calibrated hematocrit tube is spun at 1500-2000 rpm for 3 minutes.  An interface forms between the cells and the plasma at the start of the analysis.  It is tracked by the laser for the duration of the 3 minutes.  Up to 100 measurements are taken during the cycle.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | N/A |
| Specimen Collection and/or Timing | Collection tube must be filled properly and mixed thoroughly. |
| Special Collection Procedures | None defined |
| Other | N/A |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Whole Blood (K3EDTA or K2EDTA) |  |
| -Other Acceptable | None |  |
| Collection Container | 2.5 ml, 3.0 ml, 5.0 ml, or 7.0 ml Lavender top tube (K3EDTA or K2EDTA) |  |
| Volume - Optimum | 1 mL |  |
| - Minimum | 100 µL |  |
| Transport Container and Temperature | Same as above, at room temperature |  |
| Stability & Storage Requirements | Room Temperature: | (18-30C) 6 hours |
|  | Refrigerated: | (2-8C) samples should be analyzed within 12 hours of collection, bring specimen to room temperature before analysis. |
|  | Frozen: | Not appropriate |
| Timing Considerations | N/A |  |
| Unacceptable Specimens & Actions to Take | Blood collected in any tube other than EDTA; cancel test and request redraw. |  |
| Compromising Physical Characteristics | Clotted: test not performed; test cancelled Too old: test not performed entered into computer; test cancelled Frozen: test not performed; test cancelled |  |
| Other Considerations | N/A |  |

# REAGENTS

None

# CALIBRATORS/STANDARDS

Not applicable

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| SEDRite™ PLUS (2 levels) | R & D Systems, Inc, # ESRC 20002 |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | SEDRite™ PLUS (2 levels) |
| --- | --- |
| Preparation | Bring to room temperature and mix thoroughly |
| Storage/Stability | Unopened: 2° - 8° C, until expiration date Opened: 2° - 8° C, 30 days RETURN VIALS TO REFRIGERATOR WITHIN 30 MINUTES OF USE |

6.3	Frequency

Level 1 and level 2 of the SEDRite Plus should be run at least three times per day at the beginning of each shift.

6.4	Tolerance Limits

Both controls must be within the established ranges.  Each lot of control material will arrive with expected ranges.

If one or both controls are out of range, they must be repeated.

If the control is still out of range on repeat, run a fresh vial of control material.

If there continues to be a problem, notify the supervisor immediately.

No patient samples are to be reported if the controls are not in range.

6.5	Review Patient Data

Review patient results for unusual patterns, trends or distributions in patient results, such as an unusually high percentage of abnormal results.

6.6	Documentation

      - All QC results are documented in the computer.

- Refer to complete policies and procedures for QC documentation and for record retention requirements in the Laboratory QC Program.

6.7	Quality Assurance Program

- Each new lot number of reagent or new shipment of the same lot of reagent must be tested with external control materials and previously analyzed samples. Performance of the new lot must be equivalent to the previous lot.

- Training must be successfully completed and documented prior to performing this test.  This procedure must be incorporated into the departmental competency assessment program.

- The laboratory participates in CAP proficiency testing. All proficiency testing materials must be treated in the same manner as patient samples.

- Consult the Laboratory QC Program for complete details.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

ESR STAT PLUS instrument

7.2	Equipment

None

7.3	Supplies

Lithium Heparin Coated Calibrated Analysis Tubes

(kit includes:  500 analysis tubes and wipes)

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used. A current Package Insert is included in the Addendum to this procedure.

| 8.1 | Instrument Set-up Protocol |
| --- | --- |
| 1. | Insert an empty analysis tube into the “0” slot under the lip on the outside edge of the rotor. The blank should be run in the “0” slot every time the instrument is used. This blank acts as a balance. Results may be deemed to be correct even if results are determined without the index tube in place. |
| 2. | If a printer is used, be certain that the printer cable is attached to the ESR STAT PLUS, and the thermal paper roll is properly installed and the printer is turned on. |
| 3. | Turn on the ESR STAT PLUS. |
| 4. | The instrument will, upon being turned on, run through an equilibration process which will take 5 minutes. Proceed with running controls and patient samples after the instrument presents the ready screen. |

| 8.2 | Specimen / Reagent Preparation |
| --- | --- |
| 1. | The patient/control samples must be at room temperature. |
| 2. | To prepare control and patient samples, mix the whole blood samples (or controls) thoroughly. |
| 3. | If the samples are being rocked, you may pick up the EDTA blood collection tube and draw a sample immediately into the calibrated analysis tube. If the samples have not been rocked, take care to mix the sample by gentle inversion 8-10 times before drawing sample into the calibrated analysis tube. |
| 4. | When the instrument asks for the sample to be placed in the first slot, it is time to fill the calibrated analysis tube with the sample. a. remove a wipe from the wipe container. b. remove a calibrated analysis tube from the tube container. |
| 5. | Hold the calibrated analysis tube at the end with the self seal. a. remove the cap from the EDTA blood collection tube Note: face protection must be worn during this process b. tilt the blood collection tube enough to present the sample c. place the end of the calibrated analysis tube into the blood sample and draw the sample up to the fill mark d. remove from the blood sample the calibrated analysis tube when filled e. cap and place the EDTA sample in a tube rack on the counter f. with the free hand, pick up the wipe, wrap around the calibrated analysis tube up to the fill mark and draw the tube through the wipe to remove any excess blood. g. invert the calibrated analysis tube to cause the fluid to contact the self seal. |
| 6. | The sample is ready for insertion into the instrument. |

| 8.3 | Test Run |
| --- | --- |
| 1. | Open lid by pressing lid release on left side of the lid |
| 2. | Remove the rotor lock by pulling up on the lock at center of the rotor |
| 3. | PRECAUTION: Whenever you remove the plastic rotor lock, always put it on the shelf near the lid hinge. This will prevent running the analyzer without the rotor lock in place. |
| 4. | Press 1 for ESR Test |
| 5. | Insert the prepared capillary tube into slot 1 of the rotor and press 0 |
| 6. | Input the accession numbers into the identification number area |
| 7. | Press E to start |
| 8. | Fit rotor lock onto samples and close the lid, making sure lid “click” is heard |
| 9. | Press E to confirm and to begin analysis |
|  | To process more than 1 sample in a cycle: |
| 1. | After the first sample accession number is entered, press E, then press 0 |
| 2. | Input the accession number into the identification number area |
| 3. | Press E |
| 4. | Fit rotor lock onto samples and close the lid, making sure lid “click” is heard |
| 5. | Press E to confirm and to begin analysis |

At the completion of the analysis, the ESR STAT will show the following:

ESR    31     mm/hr                    press 1 to print results

ID:  79462222                          Tube: 1

(E=next, C=standby)

* If  you have run only one sample, the screen will present results and the bottom of screen will read  (C=standby)

To view the second sample result, press E=Next to view result

| 8.4 | Special Handling |
| --- | --- |
| 1. | If you are distracted and the blood sample in the blood collection tube is allowed to sit for more than one minute, be sure to re- mix the sample by gentle inversion at least eight times. |
| 2. | When not in use – The lid must be closed and locked (an audible click confirms locking of lid). |
| 3. | To retrieve prior results: Go to the Standby screen. Select C=prior results, enter the accession number for the sample result to be retrieved. The results will be presented on the screen. The ESR STAT will remember results for the last four cycles (up to 12 prior results). If the instrument is turned off, all prior results will be lost. |

Note:

If the instrument is not used within 15 minutes since the last sample is run, the instrument will automatically equilibrate for 1 minute.  For the auto equilibration to start, the instrument lid must be closed and locked.

The user may interrupt the auto equilibration by pressing STOP.   Interrupting the auto equilibration will have no adverse impact on results.

## To operate the ESR STAT PLUS after the analyzer has completed the equilibration sequence; the standby screen will automatically appear as follows:

## ESR Test

- Date-Time Settings

(C=previous results)

# CALCULATIONS

None

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

None required

10.2	Rounding

N/A

10.3	Units of Measure

mm/hr

10.4	Clinically Reportable Range (CRR)

N/A

  - Repeat Criteria and Resulting

| IF the result is … | THEN… |
| --- | --- |
| Outside of the AMR | Repeat testing, ensuring that the sample is mixed well. |

# EXPECTED VALUES

11.1	Reference Ranges

**Reference rages pending validation.**

11.2	Priority 1 & 2 Limits

None established

11.3	Priority 3 Limit(s)

None established

# CLINICAL SIGNIFICANCE

The Erythrocyte Sedimentation Rate is a widely used, non-specific screening test.  It is indicative of the presence of infectious, inflammatory, degenerative or neoplastic conditions.  The increased rate of red cell determination is mainly associated with qualitative and quantitative changes in the plasma proteins.

# PROCEDURE NOTES

- FDA Status:   FDA Approved/cleared

- Validated Test Modifications:  None

- The patient sample must be filled to the fill line properly, not more than +/- 1 mm.

- The patient/control sample must be at room temperature

- The capillary tube should not lay down in a horizontal position for more than 5 minutes

- The analysis tube must be wiped clean of any blood before starting analysis

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

1 - 145 mm/hr

14.2	Precision

N/A

14.3	Interfering Substances

N/A

14.4	Clinical Sensitivity/Specificity/Predictive Values

N/A

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- If you drive on company business (in a company car, a rental, or your own personal vehicle) you must comply with the requirements in the "Driver Safety Manual".  Please remember that driving requires your FULL attention, and multi-tasking while driving (e.g., talking on a cell phone) is an unacceptable risk that you must not take.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

- Laboratory Safety Manual

- Material Safety Data Sheets (MSDS)

- Laboratory QC Program

- Repeat Testing Requirements (Laboratory policy)

# REFERENCES

- HemaTechnologies, Inc. User Manual ver. 3.5

- ERYTHROCYTE SEDIMENTATION RATE (ESR-W) WESTERGREN PROCEDURE 1998

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |

# ADDENDA

| Addendum | Title |
| --- | --- |